## Slide 1: BPM System Design
Invited tutorial at IBIC 2015, Melbourne
Hermann Schmickler, CERN
With most slides from:
Peter Forck (GSI), Manfred Wendt (CERN), Piotr Kowina (GSI), Rhodri Jones (CERN), Marek Gasior (CERN), T.Lefevre (CERN) David Rubin (Cornell), M.Tejima (KEK)

## Slide 2: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook
Not covered: Simulation Tools, Managerial Considerations
Speed

## Slide 3: Introduction
The particle beam induces an electrical signal in two opposing electrodes. The induced signals depend - on the beam intensity - on the proximity of the beam to the sensors
Intuitively the term: “difference over sum” should measure the beam position….how well we will see later
One of the principle problems is that the requested observable (beam position) is measured as the difference of two large numbers…. We will see what impact this has on the measurement resolution/accuracy/precision
Some smart people build monitors, which give the difference by the nature of the monitor…cavity BPMs. Those we will treat separately. In general they are limited to applications in linear accelerators due to their high coupling impedance.

## Slide 4: Application of BPMs
On bunched beam
Trajectory Measurement: Measure beam positions during one revolution/passage through transfer line
Closed Orbit Measurement (CO): Average over “many” trajectories
Time resolution: from long averages of CO up to turn by turn trajectories, turn by turn trajectories even bunch by bunch  depending on bunch length even observation of bunch shape
Derived quantities: - tune, chromaticity, coupling (using excited betatron oscillations plus observation with a single BPM) - β-function and phase advance around the ring (using excited betatron oscillations plus observation with all BPMs), dispersion, injection matching…
Usage in real time feedbacks (on CO, multibunch stability)
Sensitivity down to nA beams (pC)

## Slide 5: Accuracy, Precision, Resolution
Very often confused in day-to-day language
Accuracy:= also called trueness of measurement
Precision:= indicates how well one can reproduce measurements
Resolution:= smallest possible difference in successive measurements
Ex: BPM: Mechanical and electrical offsets, gain factors influence the accuracy, various noise sources or timing jitter influence the precision, quantization in the ADC can limit the resolution.

## Slide 6: Functional Specification (FS)
Make sure you have this (in writing) before you start developing. Initiative normally comes from the Instrument responsible, FS has to be written by future users of instrument.
FS needs to contain - beam parameters, - modes of operation, - required accuracy, precision, resolution - expected frequency of usage (i.e. wire scanners) - specifications for control software and data analysis tools
Leads to an Engineering Specification (ES) produced by instrument specialist.
ES and FS represent a “contract” between producer and user.
Examples: http://sl-div-bi.web.cern.ch/sl-div-bi/LHC/ParamAndLayouts/Doc/FuncSpec.htm

## Slide 7: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals: capacitive sensors
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook

## Slide 8
General Idea: Detection of Wall Charges
The image current at the vacuum wall is monitored on a high frequency basis
i.e. the ac-part given by the bunched beam.
For relativistic velocities, the electric field is mainly transversal:
Beam Position Monitor BPM equals Pick-Up PU

## Slide 9
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
Principle of Signal Generation of capacitive BPMs
Animation by Rhodri Jones (CERN)
Same behavior as top plate

## Slide 10
Model for Signal Treatment of capacitive BPMs
The wall current is monitored by a plate or ring inserted in the beam pipe:
The image current Iim at the plate is given by the beam current and geometry:
Using a relation for Fourier transformation: Ibeam = I0eiωt ⇒ dIbeam/dt = iωIbeam.

## Slide 11
Example of Transfer Impedance for Proton Synchrotron
The high-pass characteristic for typical synchrotron BPM:
Uim(ω)=Zt(ω)·Ibeam(ω)
Parameter for shoe-box BPM:
C=100pF, l=10cm, β=50%
fcut= ω/2π=(2πRC)-1
for R=50   fcut= 32 MHz
for R=1 M  fcut= 1.6 kHz

## Slide 12
Signal Shape for capacitive BPMs: differentiated ↔ proportional
Depending on the frequency range and termination the signal looks different:
High frequency range ω >> ωcut :
 direct image of the bunch. Signal strength Zt  A/C i.e. nearly independent on length
Low frequency range ω << ωcut:
 derivative of bunch, single strength Zt  A, i.e. (nearly) independent on C
Example from synchrotron BPM with 50  termination (reality at p-synchrotron : σ>>1 ns):

## Slide 13
Calculation of Signal Shape: Bunch Train
Train of bunches with R=50  termination  f << fcut :
Parameter: R=50   fcut=32 MHz, all buckets filled
C=100pF, l=10cm, β=50%, σt=100 ns
Calculation:
Fourier spectrum is composed of lines separated by acceleration frf
Envelope given by single bunch Fourier transformation
Differenciated bunch shape due to fcut >> frf
Remark: 1 MHz< frf <10MHz  Bandwidth 100MHz = 10∙frf for broadband observation.

## Slide 14
Shoe-box BPM for Proton or Ion Synchrotron
Frequency range: 1 MHz < frf < 10 MHz  bunch-length >> BPM length.
Signal is proportional to actual plate length:
In ideal case: linear reading
Shoe-box BPM:
Advantage: Very linear, low frequency dependence
i.e. position sensitivity S is constant
Disadvantage: Large size, complex mechanics
high capacitance
Size: 200x70 mm2

## Slide 15
Technical Realization of Shoe-Box BPM
Quadrupole
Linear Ampl.
Beam
70
180
Linear Ampl.
Right
channel
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.

## Slide 16
70
180
Linear Ampl.
Right
channel
Technical Realization of Shoe-Box BPM
Technical realization at HIT synchrotron of 46 m length for 7 MeV/u440 MeV/u
BPM clearance: 180x70 mm2, standard beam pipe diameter: 200 mm.

## Slide 17: Other Types of diagonal-cut BPM
Round type: cut cylinder
Same properties as shoe-box:
Wound strips:
Same distance from beam and capacitance for all plates
But horizontal-vertical coupling.
Vertical BPM
Horizontal BPM
Guard Ring
Signal Out
Other realization: Full metal plates
No guard rings required
but mechanical alignment more difficult

## Slide 18
Button BPM Realization
LINACs, e--synchrotrons: 100 MHz < frf < 3 GHz  bunch length  BPM length
 50  signal path to prevent reflections
Button BPM with 50  
Example: LHC-type inside cryostat:
24 mm, half aperture a =25 mm, C =8 pF
 fcut= 400 MHz, Zt = 1.3  above fcut
Ø24 mm
From C. Boccard (CERN)

## Slide 19
2-dim Model for Button BPM
‘Proximity effect’: larger signal for closer plate
Ideal 2-dim model: Cylindrical pipe  image current density
via ‘image charge method’ for ‘pencil’ beam:
Image current: Integration of finite BPM size:

## Slide 20: 2-dim Model for Button BPM
In addition, frequency dependence can be calculated by analytic model or numerically.
button
Ideal 2-dim model: Non-linear behavior and hor-vert coupling:
Sensitivity: x=1/S∙ΔU/ΣU with S [%/mm] or [dB/mm]
For this example: center part S=7.4%/mm  k=1/S=14mm
The measurement of U delivers:  here Sx = Sx(x, y) i.e. non-linear.

## Slide 21
Button BPM at Synchrotron Light Sources
Beam position swept with 2 mm steps
Non-linear sensitivity and hor.-vert. coupling
At center Sx = 8.5%/mm in this case
From S. Varnasseri, SESAME, DIPAC 2005
Due to synchrotron radiation, the button insulation might be destroyed
buttons only in vertical plane possible  increased non-linearity
Optimization: horizontal distance and size of buttons

## Slide 22
Comparison Shoe-Box and Button BPM
 | Shoe-Box BPM | Button BPM
Precaution | Bunches longer than BPM | Bunch length comparable to BPM
BPM length (typical) | 10 to 20 cm length per plane | 1 to 5 cm per button
Shape | Rectangular or cut cylinder | Orthogonal or planar orientation
Bandwidth (typical) | 0.1 to 100 MHz | 100 MHz to 5 GHz
Coupling | 1 M or 1 k (transformer) | 50 
Cutoff frequency (typical) | 0.01… 10 MHz (C=30…100pF) | 0.3… 1 GHz (C=2…10pF)
Linearity | Very good, no x-y coupling | Non-linear, x-y coupling
Sensitivity | Good, care: plate cross talk | Good, care: signal matching
Usage | At proton synchrotrons, frf < 10 MHz | All electron acc., proton Linacs, frf > 100 MHz

## Slide 23: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals: stripline BPMs
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook

## Slide 24
Stripline BPM: General Idea
For short bunches, the capacitive button deforms the signal
Relativistic beam β1  field of bunches nearly TEM wave
Bunch’s electro-magnetic field induces a traveling pulse at the strips
Assumption: Bunch shorter than BPM, Zstrip=R1=R2=50  and vbeam=cstrip.
LHC stripline BPM, l=12 cm
From C. Boccard, CERN

## Slide 25
Stripline BPM: General Idea
For relativistic beam with β1 and short bunches:
Bunch’s electro-magnetic field induces a traveling pulse at the strip
Assumption: lbunch<< l, Zstrip=R1=R2=50  and vbeam=cstrip
Signal treatment at upstream port 1:
Signal at downstream port 2: Beam induced charges cancels with traveling charge from port 1
 Signal depends on direction  directional coupler: e.g. can distinguish between e- and e+ in collider
t=0: Beam induced charges at port 1:
 half to R1, half toward port 2
t=l/c: Beam induced charges at port 2:
 half to R2, but due to different sign,
it cancels with the signal from port 1
 half signal reflected
t=2∙l/c: reflected signal reaches port 1
If beam repetition time equals 2∙l/c: reflected preceding port 2 signal cancels the new one:
 no net signal at port 1

## Slide 26
Stripline BPM: Transfer Impedance
The signal from port 1 and the reflection from port 2 can cancel  minima in Zt .
For short bunches Ibeam(t)  Ne ∙ (t):
Zt show maximum at l=c/4f=λ/4 i.e. ‘quarter wave coupler’ for bunch train
 l has to be matched to vbeam
No signal for l=c/2f=λ/2 i.e. destructive interference with subsequent bunch
Around maximum of |Zt|: phase shift φ=0 i.e. direct image of bunch
fcenter=1/4 ∙ c/l ∙ (2n-1). For first lope: flow=1/2∙fcenter, fhigh=3/2 ∙ fcenter i.e. bandwidth 1/2∙fcenter
Precise matching at feed-through required t o preserve 50  matching.
Stripline length l=30 cm, α=100

## Slide 27
Stripline BPM: Finite Bunch Length
The signal at port 1 for a finite bunch of length σ: .
Zt(ω) decreases for higher frequencies
If total bunch is too long (±3σt>l) destructive interference leads to signal damping
Cure: length of stripline has to be matched to bunch length
Stripline length l=30 cm, α=100
Caution: Zt depends on beam’s bunch length σ

## Slide 28
Realization of Stripline BPM
20 cm stripline BPM at TTF2 (chamber 34mm)
And 12 cm LHC type:
From . S. Wilkins, D. Nölle (DESY), C. Boccard (CERN)

## Slide 29: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals: stripline-BPMs
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook

## Slide 30
Cavity BPM: Principle
High resolution on t < 1 μs time scale can be achieved by excitation of a dipole mode:
TM01
TM11
TM02
f / GHz
U / V
TM010
TM011
TM020
U~Q
U~Qr
U~Q
Application:
small e- beams
and short pulses < ns
(ILC, X-FEL…)
‘δ-excitation’
→oscillation with
Q  1000 and τ =2Q/2f  100 ns
From D. Lipka,
DESY, Hamburg
For pill box the resonator modes given by geometry:
monopole TM010 with f010
 maximum at beam center  strong excitation
Dipole mode TM011 with f011
 minimum at center  excitation by beam offset
Detection of dipole mode amplitude

## Slide 31
Cavity BPM: Example of Realization
From M. Wendt (FNAL)
FNAL realization:
Cavity:  113 mm
Gap 15 mm
Mono. f010=1.1GHz
Dipole. f110=1.5GHz
Qload  600
With comparable BPM
0.1 μm resolution
within 1 μs
Basic consideration for detection of Eigen-frequency amplitudes:
Dipole mode f110 separated from monopole mode
but to finite quality factor Q  Δf=f/Q
Frequency f110  1…10 GHz
Waveguide house the antennas
  (task: suppression of TM010 mode signal)
Gap
15 mm
Antenna for
monopole mode
Antenna for
dipole mode
Waveguide input for dipole mode
Cavity Ø113mm
Antenna for
dipole mode
Beam pipe
Ø39mm
f010
f110
frequency
Ampl.

## Slide 32
Suppression of mono-pole mode: waveguide that couple only to dipole-mode
due to fmono < fcut < fdipole
Monopole Mode
Dipole Mode
Courtesy of D. Lipka,
DESY, Hamburg
Prototype BPM for ILC Final Focus:
Required resolution of 5 nm (yes nano!) in a 6×12 mm diameter beam pipe
Achieved world record resolution of 8.7 nm 0.28(stat) 0.35(sys) nm
at ATF2 (KEK, Japan).
Courtesy of D. Lipka,
DESY, Hamburg
wave guide
Mono-pole mode
Dipole-pole mode
Courtesy of D. Lipka & Y. Honda
Courtesy of D. Lipka and Y. Honda
Cavity BPM: Suppression of monopole Mode
Notes: 32

## Slide 33
Comparison of BPM Types (simplified)
Remark: Other types are also some time used, e.g. wall current, inductive antenna,
BPMs with external resonator, slotted wave-guides for stochastic cooling etc.
Type | Usage | Precaution | Advantage | Disadvantage
Shoe-box | p-Synch. | Long bunches frf < 10 MHz | Very linear No x-y coupling Sensitive For broad beams | Complex mechanics Capacitive coupling between plates
Button | p-Linacs, all e- acc. | frf > 10 MHz | Simple mechanics | Non-linear, x-y coupling Possible signal deformation
Stipline | colliders p-Linacs all e- acc. | best for β1, short bunches | Directivity ‘Clean’ signal Large signal | Complex 50 Ω matching Complex mechanics
Cavity | e- Linacs (e.g. FEL) | Short bunches Special appl. | Very sensitive | Very complex, high frequency

## Slide 34: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals: capacitive sensors
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook

## Slide 35: The Ideal BPM Read-out Electronics!?
Time multiplexing of the BPM electrode signals:
  Interleaving BPM electrode signals by different cable delays
  Requires only a single read-out channel!
Beam
ADC
ADC
ADC
ADC
FPGA
Fiber Link
DAQ
PS
CLK
BPM pickup (e.g. button, stripline)
Digital BPM electronics (rad-hard, of course!)
Very short coaxial cables
“Super” ADCs
“Monster” FPGA
“Ultra” low jitter clock!

## Slide 36: BPM Building Blocks
BPM pickup
  RF device, EM field detection, center of charge
  Symmetrically arranged electrodes, or resonant structure
Read-out electronics
  Analog signal conditioning
  Signal sampling (ADC)
  Digital signal processing
Analog Signal Conditioning
Digital Signal Processing
Data Acquisition
Trigger & Timing
Control
Power Supply & Misc.
BPM Pickup
position
data
control
system
(LAN)
timing,
trigger
signals
feedback bus
(if applicable)
  Data acquisition and control system interface
  Trigger, CLK & timing signals
  Provides calibration signals or other drift compensation methods
CAL
ADC
CLK
Minimize?!

## Slide 37: Signal Processing & Normalization
Extract the beam position information from the electrode signals: Normalization
  Analog using Δ-Σ or 900-hybrids, followed by filters, amplifiers mixers and other elements, or logarithmic amplifiers.
  Digital, performing the math on individual digitized electrode signals.
Decimation / processing of broadband signals
  BPM data often is not required on a bunch-by-bunch basis
    Exception: Fast feedback processors
    Default: Turn-by-turn and “narrowband” beam positions
  Filters, amplifiers, mixers and demodulators in analog and digital to decimate broadband signals to the necessary level.
Other aspects
  Generate calibration / test signals
  Correct for non-linearities of the beam position response of the BPM
  Synchronization of turn-by-turn data
  Optimization on the BPM system level to minimize cable expenses.
  BPM signals keep other very useful information other than that based on the beam displacement, e.g.
    Beam intensity, beam phase (timing)

## Slide 38: For the “Oldies”: Analog Signal Processing Options
D I G I T I Z E R
turn by turn
no turn by turn
Legend:
/ Single channel
Wide Band
Narrow band
Normalizer
Processor
Active
Circuitry
Heterodyne
POS = (A-B)
Synchronous
Detection
AGC
on S
MPX
Electrodes
A, B
Passive
Normaliz.
POS = [log(A/B)]
= [log(A)-log(B)]
Differential
Amplifier
Logarithm.
Amplifiers
Individual
Treatment
Limiter,
Dt to Ampl.
Amplitude
to Time
POS = [A/B]
POS = [ATN(A/B)]
Amplitude
to Phase
.
Limiter,
f to Ampl.
POS = D / S
Heterodyne
Hybrid
D / S
Homodyne
Detection
POS = D / S or
= (A-B)/(A+B)
Sample,Track,
Integr. & Hold
Switch. gain
Amplifier
courtesy G. Vismara (BIW 2001)

## Slide 39: New member of analog orbit measurements:
Further development of “BBQ”-tune measurement system (M.Gasior, CERN)

## Slide 40: Compensated Diode Detector for BOM
Sub-micrometre resolution can be achieved with relatively simple hardware and signals from any position pick-up.
To be used for the future LHC collimators with embedded BPMs.
courtesy M. Gasior

## Slide 41: Digital BPM Signal Processing
Why digital signal processing?
  Better reproducibility of the beam position measurement
    Robust to environmental conditions, e.g. temperature, humidity, (radiation?)
    No slow aging and/or drift effects of components
    Deterministic, no noise or statistical effects on the position information
  Flexibility
    Modification of FPGA firmware, control registers or DAQ software to adapt to different beam conditions or operation requirements
  Performance
    Often better performance, e.g. higher resolution and stability compared to analog solutions
    No analog equivalent of digital filters and signal processing elements.
BUT: Digital is not automatically better than analog!
  Latency of pipeline ADCs (FB applications)
  Quantization and CLK jitter effects, dynamic range & bandwidth limits
  Digital BPM solutions tend to be much more complex than some analog signal processing BPM systems
    Manpower, costs, development time

## Slide 42: BPM Read-out Electronics
ADC
900
CIC
FIR
Σ
MEMORY
NCO
I-Channel
Q-Channel same as I
NB
WB
raw
A
B
C
D
BPF
Att
Coordinate
Transformation
A-Electrode Analog Conditioning
B, C, D Analog same as A
Ctrl
CLK & Timing
A Data
ADC
900
CIC
FIR
Σ
MEMORY
NCO
I-Channel
Q-Channel same as I
NB
WB
raw
A
B
C
D
BPF
Att
BPF
LPF
Coordinate
Transformation
A-Electrode Analog Conditioning
B, C, D Analog same as A
Ctrl
LO
CLK & Timing
A Data
Do we need this analog section?

## Slide 43: “Ringing” Bandpass-Filter (BPF)
Single Bunch Responses:
Stripline BPM: Bessel
Button BPM: Bessel
Stripline BPM: Butterworth
fcenter: 500 MHz BW: 25 MHz
BPM electrode signal energy is highly time compressed
  Most of the time: “0 volt”!
A “ringing” bandpass filter “stretches” the signal
  Passive RF BPF
    Matched pairs!
  fcenter matched to frev or fbunch
    Quasi sinusoidal waveform
  Reduces output signal level
    Narrow BW: longer ringing, lower signal level
  Linear group delay designs
    Minimize envelope ringing
    Bessel, Gaussian, time domain designs

## Slide 44: “Ringing” BPF & Multi-Bunches
Bunch spacing < BPF ringing time:
  Superposition of single bunch BPF responses
  More continuous “ringing”, smearing of SB responses
Bunch spacing < BPF rise time
  Constructive signal pile-up effect
    Output signal level increases linear with decreasing bunch spacing

## Slide 45: Analog Digital Converter
Quantization of the continuous input waveform at equidistant spaced time samples
  Digital data is discrete in amplitude and time
LSB voltage (resolution)
  E.g. 61 μV (14-bit), 15 μV (16-bit) @ 1 volt VFSR
Quantization error (dynamic range)
  E.g. 84 dB (14-bit), 96 dB (16 bit)
SNR limit due to aperture jitter
  E.g. 62 dB@500 MHz, 0.25 psec (equivalent to EOB=10.3)
M=3-bit ADC

## Slide 46: ADC Technology
 | Type | Res. [bit] | Ch. | Power [W] | fs (max) [MSPS] | BW [MHz] | SNR @ fin [dB @ MHz]
AD | AD9652 | 16 | 2 | 2.2 | 310 | 485 | 72 @170
AD | AD9680 | 14 | 2 | 3.3 | 1000 | 2000 | 67 @ 170
LT | LTM9013* | 14 | 2 | 2.6 | 310 | 300* | 62 @ 150
TI | ADC16DX370 | 16 | 2 | 1.8 | 370 | 800 | 69 @ 150
TI | ADS5474-SP | 14 | 1 | 2.5 | 400 | 1280 | 70 @ 230
* has an analog I-Q mixer integrated, 0.7 GHz < fin< 4 GHz
Dual Channel
  I-Q sampling with separate ADCs
Pipeline architecture
  Continuous CLK
  Data latency
A-D mixed designs
  Mixers, gain, filters, etc.

## Slide 47: Sampling Theory
A band limited signal x(t) with B=fmax can be reconstructed if
  Nyquist-Shannon theorem
  The exact reconstruction of x(t) by xn=x(nT):
Aliasing of a sampled sin-function
  Samples can be interpreted by
to be avoided!
courtesy Wikipedia

## Slide 48: I-Q Sampling
Vector representation of sinusoidal signals:
  Phasor rotating counter-clockwise (pos. freq.)
Q
I
A
I
Q
φ0
I: in-phase component
Q: quadrature-phase component
I-Q sampling at:
Q(t0)
I(t1)
A=1.33

## Slide 49: I-Q Demodulation of BPM Signals
Digitized sinusoidal waveform data is sampled, or undersampled at fs/n = 4xfin
  Also called digital down-converter (DDC)
  Gives amplitude and phase of the input signal
    Phase reference could be fs (CLK), of a separate reference signal
n=2

## Slide 50: Digital Down-Converter
Goals
  Convert the band limited RF-signal to baseband (demodulation)
  Data reduction (decimation)
DDC Building blocks:
  ADC
    Single fast ADC (oversampling)
  Local oscillator
    Numerically controlled oscillator (NCO) based on a direct digital frequency synthesizer (DDS)
  Digital mixers (“ideal” multipliers)
  Decimating low pass (anti alias) filters
    Filtering and data decimation.
    Implemented as CIC and/or FIR filters
courtesy T. Schilcher

## Slide 51: Signal/Noise & Theoretical Resolution Limit
Minimum noise voltage at the 1st gain stage:
  With the stripline BPM and Bessel BPF example: R = 50 Ω, Δf = 25 MHz  vnoise = 4.55 μV (-93.83 dBm)
Signal-to-noise ratio:
  Where Δv is the change of the voltage signal at the 1st gain stage due to the change of the beam position (Δx, Δy).
  Consider a signal level v ≈ 22.3 mV (-20 dBm)
    Bessel BPF output signal of the stripline BPM example
  22.3 mV / 4.55 μV ≈ 4900 (73.8 dB) would be the required dynamic range to resolve the theoretical resolution limit of the BPM
    Under the given beam conditions, e.g. n=1e10, σ=25mm, single bunch, etc.
    The equivalent BPM resolution limit would be: Δx=Δy=0.66μm (assuming a sensitivity of ~2.7dB/mm)

## Slide 52: S/N & BPM Resolution (cont.)
Factors which reduce the S/N
  Insertion losses of cables, connectors, filters, couplers, etc.
    Typically sum to 3…6 dB
  Noise figure of the 1st amplifier, typically 1…2 dB
  The usable S/N needs to be >0 dB, e.g. 2.3 dB is sometimes used as lowest limit. (HP SA definition)
  For the given example the single bunch / single turn resolution limit reduces by ~10 dB (~3x): 2…3 μm
Factors to improve the BPM resolution
  Increase the signal level
    Increase BPM electrode-to-beam coupling, e.g. larger electrodes
    Higher beam intensity
  Increase the measurement time, apply statistics
    Reduce the filter bandwidth (S/N improves with 1/√BW)
    Increase the number of samples (S/N improves with √n)

## Slide 53: BPM Read-out Electronics
Typical BPM read-out scheme
  Pipeline ADC & FPGA
    14-16 bit, >300 MSPS, ~70 dB S/N
  Separate analog signal processing for the channels
A
B
C
D
BPF
Att
A-Electrode Analog Conditioning
B, C, D Analog same as A
Ctrl
ADC
900
CIC
FIR
Σ
MEMORY
NCO
I-Channel
Q-Channel same as I
NB
WB
raw
A
B
C
D
BPF
Att
Coordinate
Transformation
A-Electrode Analog Conditioning
B, C, D Analog same as A
Ctrl
LO
CLK & Timing
A Data
Choices:
  Analog downconverter?!
  RF locked (sync) CLK & LO signals?!
    No I-Q required
A
B
C
D
BPF
Att
A-Electrode Analog Conditioning
B, C, D Analog same as A
Ctrl
with analog downconverter
ADC
900
CIC
FIR
Σ
MEMORY
NCO
I-Channel
Q-Channel same as I
NB
WB
raw
A
B
C
D
BPF
Att
BPF
LPF
Coordinate
Transformation
A-Electrode Analog Conditioning
B, C, D Analog same as A
Ctrl
LO
CLK & Timing
A Data

## Slide 54: Analog Downconverter vs. Direct Digital Under-Sampling
ADC dynamic range is limited to ~70 dB
  Not sufficient for most BPM applications
  Need for analog signal attenuator / gain stages
    Requires calibration signal to avoid “electronic offsets”
Analog downconverter
  Certainly necessary for the conditioning of cavity BPM signals
  Allows sampling in the 1st Nyquist passband (no undersampling)
  Relaxes input RF filter requirements
  Relaxed ADC and CLK requirements
  May relax cable requirements and improve S/N
    Analog hardware installation near the BPM pickup
    Transfer analog IF signals out of the tunnel
  Additional analog hardware required
  Generates additional image frequencies
    Consider image rejection analog mixer!

## Slide 55: Example: ATF DR BPM Signal Processing
ADC Input
14 Bits
71.4 MHz
NB Filter
1.4kHz output
16 Bits/ch
32 ch
/
NB Gate
WB Gate(s)
DDR RAM
NB Data
TBT Data
Raw Data
Σ
50Hz
VME
NB Data
VME
Raw Data
TBT Filter
VME
TBT Data
8 ch
/
Trigger
DAQ SM
Ch delays (clocks)
Gates in Turns
WB Gate(s)
NB Gate
32 Registers
VME
NB Sums
VME
IRQ
reset
latch
reset
latch
8 ch
/
average NB data with n/fpower

## Slide 56: ATF BPM Narrowband Signal Processing
Process 8 ADC channels in parallel up to FIR filter
  Digitally downconvert each channel into I,Q then filter I,Q independently
  CIC Filters operating in parallel at 71.4MHz
    Decimate by 17KSPS to 4.2KSPS output rate
  1 Serial FIR Filter processes all 32 CIC Filter outputs
    80 tap FIR (400 Hz BW, 500 Hz Stop, -100 db stopband) -> 1KHz effective BW
    Decimate by 3 to 1.4 KSPS output rate -> ability to easily filter 50Hz
  Calculate Magnitude from I,Q at 1.4KHz
    Both Magnitude and I,Q are written to RAM
    Also able to write I,Q output from CIC to RAM upon request
ADC Input
14 Bits
69 MHz
X
NCO (sin, cos)
24 Bits Phase
(~1 Hz)
I
Q
16 Bits
CIC
5 Stages
R=17001
DDC
24 Bits
4.2 KSPS
FIR (80 taps)
LPF 500Hz
Decimate 3
Bit
Shift
Select
Significant
Bits
20 Bits
4.2 KSPS
16 Bits
1.4 KSPS
I
Q
- Denotes Peak Detectors to optimize scaling

## Slide 57: ATF DR Turn-by-Turn Beam Studies
Beam optics studies with 96 BPMs in the ATF damping ring
β function measurement
β beating measurement
ϕ measurement
courtesy Y. Renier

## Slide 58: Libera BPM Electronics
Analog & digital crossbar switch
  Compensation of long term drift effects in the analog sections
courtesy Instrumentation Technologies

## Slide 59: Libera BPM Performance
< 0.5 µm RMS at turn-by-turn data rate
40 nm RMS at 10 kS/s data rate (0.01 – 1 kHz)
10 nm RMS for slow monitoring
sub-micron longterm stability
Temperature drift < 200 nm / °C
Full Fast Orbit Feedback implementation with magnet output
Fast Interlock detection (< 100 µs)
Clean turn to turn measurement using Time-Domain Processing
Libera Brilliance +
Electron beam position measurements
Beam Current Dependence
courtesy P. Leban

## Slide 60: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals: capacitive sensors
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook

## Slide 61: Synchronization, accelerator timing
Needs in itself a complete talk…
Three major tasks: 1) clocking of front-end electronics  distribution of clock reference via copper or fibre links continuous synchronization of local PLL circuits for clocking 2) fast accelerator related timing  identification of bunches around machine 3) slow accelerator related timing  indicates machine state transitions (injection…acceleration, dump)
Typically provided over one link as encoded messages, which are modulated onto a carrier frequency (=clock)
In most labs this service is provided by the controls team
Lesson: Always have an acquisition mode, which does not depend on the timing system (self trigger)…most essential during commissioning

## Slide 62: Beam Threading
Threading the beam round the LHC ring (very first commissioning in 2008)
  One beam at a time, one hour per beam.
  Collimators were used to intercept the beam (1 bunch, 2×109 protons)
  Beam through 1 sector (1/8 ring)
    correct trajectory, open collimator and move on.
Beam 2 threading
BPM availability ~ 99%
Courtesy of CMS
Courtesy of ATLAS

## Slide 63: Outline
Principle, Applications of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals: capacitive sensors
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook

## Slide 64: From Precision to Accuracy
BPM sensor: - matching of sensor pairs - mechanical pre-alignment to nearest quadrupole magnetic axis (~100 um); further refinement of offsets with k-modulation of quadrupole - linearization of response by post-processing based on characterization of BPM response on a measurement bench - thermal stabilization of BPM body (= one of the reasons for top-up injection)
BPM electronics (prior to beam) - measurement of insertion losses and phase errors of whole RF signal path - calculation/measurement of coupling impedance to beam - measurement of gain factors of each acquisition chain
Various setups in order to inject calibration tones as close to the sensor as possible -> button systems: inject in one button + capacitive coupling to 2 adjacent buttons -> stripline couplers: use 2nd unused port for calibration tone injection -> use coupling transformers in RF-frontend
Thermal stabilization of (analog) readout electronics
Crossbar switches in the RF front-end: Periodically reassign input signals to the different acquisition chains. Deduce gain factors from signal changes  show a few examples (also taken from this conference):

## Slide 65: Long-Term Drift Compensation (as used in MANY LIGHTSOURCES)
Libera crossbar switching technique
  <100 nm stability over 14 hours
Calibration tone technique (only in narrowband operation) ATF (KEK)
courtesy N. Eddy
courtesy P. Leban

## Slide 66: STEP 1, Mapping of BPM head
Test bench for the mapping
result for mapping
All BPMs were mapped at a test bench with a movable antenna .
LER
HER

## Slide 67
Button BPM at Synchrotron Light Sources
From S. Varnasseri, SESAME, DIPAC 2005
dAB=24 mm
dAB=24 mm
vertical
horizontal
dAB=24 mm
dAB
2-dim electro-static simulation:
Beam position swept with 2 mm steps
2 mm distance between dots
Result:
Hori. Sx= 8.5%/mm
Vert. Sy = 5.6%/mm
x&y dependent polynomial fit possible

## Slide 68: STEP 2, Alignment of BPM heads against to the Q-magnets
Photograph of Measurement tool
Results of alignment of BPM heads against to the reference plane of the Q-Magnets
Measurement of the mechanical offsets of the BPM heads to the Q-magnets.

## Slide 69: STEP-3, Gain calibration of the electronics Attenuation of cable, switch, electronics, etc.
We measured the distribution of signal attenuation of the all electronics
We used a dummy head instead of BPM heads.
Results for ration between output signal B,C and D against to A in all electronics
A
B
C
D
Output signals: A, B, C, D

## Slide 70: Beam based calibration
Beam Based Alignment
Measurement of the offset of BPM to the field center of the adjacent Q-magnet using the beam.
Beam Based Gain Calibration
Calibration of the gain imbalance among four outputs of a BPM using the beam.

## Slide 71: Beam based alignment(BBA) - Principle -
BBA is searching for the beam orbit which is insensitive to the change of field strength in the quadrupole = magnetic axis of this quad.
The measured beam position for this orbit corresponds to origin offset of the BPM.
The offset gets introduced into correction tables for the BPM readings.
DXm=0 then dxm offset

## Slide 72: Actual procedure for BBA
X=0.4913[mm]
Offset
Bump orbits: -8, 0, 8[mm]
IQ : -4, -2, 0, 2, 4 [A]
-current of correction coil
from figure (a),(b),(c),
Dx/DIQ and x (IQ=0),
where x=xm
Alternative method (sinusoidal modulation of q-current at low frequencies) + synchronous detection of orbit variations at the modulation frequency….

## Slide 73
Gain mapping
Button to button gain errors
Any difference in the effective
button gains, electronic or physical,
couple real horizontal dispersion
into measured vertical dispersion
(and the horizontal dispersion is large)
Slides provided by: D. Rubin, Cornell University
8-September 2015

## Slide 74
Signal at each button depends on bunch current (k) and position (x,y)
Signals on the four buttons are related by symmetry
Combining sums and differences we find the following relationship, good to second order
Characterization of BPM Gain Errors
B3
B4
B1
B2

## Slide 75
The gains are the set of that minimize
Gain mapping with turn by turn data
A single beam passage gives a measurement of intensity at each of four button electrodes
N beam passages gives

## Slide 76
Turn by turn data
1024 measurements
Fitted gains g1-4 = (1.007,1.016,0.9434,1.035)

## Slide 77: Outline
Principle, Usage of BPMs
Functional Specifications (FS)
Sensors + Sensor Signals
Electronics
Synchronization, accelerator timing
From precision to accuracy
Outlook - new LHC orbit system for LHC-LS3 (2024) - Electro-optical BPMs  Intrabunch Measurements for <ns bunches

## Slide 78: LHC BPM Timing Specs – 1st thoughts
Upgrade of the LHC BPM electronics (for HL-LHC):
  Keep existing infrastructure
    Beam pickups, cables, fibers, VME system, etc.
  Upgrade of analog, DAQ and trigger electronics
    Minimalistic analog front-end
      Preferable time-multiplexed single channel for 2 BPM electrodes
    Digital direct downconversion (DDC) technique
      Convolution integral or I/Q mixing?
    Internal ADC and FPGA clocks with ultra-low sub-ps jitter!
    External RF bucket tagging trigger!
Relaxed external timing requirements:
  RF bucket synchronous trigger signals in 2.5 ns increments
    Jitter & drift effects <<1 ns
    Needs to be supplied to each BPM DAQ electronics
      Of course separate trigger signals for B1 and B2

## Slide 79: Conceptual Idea
A
B
delay line
BPF
Att
Ctrl
CLK & Timing
single or
dual (I/Q)
analog signal conditioning
beam pickup
ADC
CIC
FIR
CIC
FIR
LNA
convolution based DDC
or I/Q mixer
filtering & decimation
BPF response waveform
or NCO
digital direct downconversion: - convolution integral
- I/Q mixing
Brute-force digitalization
  E.g. TI ADC12J4000
    12-bit, 4 GSPS, SFDR ≈ 67 dBFS, SINAD ≈ 55 dBFS (@350MHz)
    Intergraded DDC, decimation by 4 or 8 (400MHz pass-band)
      SFDR ≈ 75 dBFS, SINAD ≈ 63 dBFS (@600MHz)
  Internal CLKs for ADC and FPGA
  External 40 MHz trigger (2.5ns steps) for bunch tagging (memory)
Single analog processing channel
  Based on delay-line networks and band-pass filters

## Slide 80: Analog Signals
Gaussian bunch:
  n=1E10, σ=25mm, v=c0
  vertical offset=1mm
BPM sensitivity:
  e.g.: 2.7 dB / mm
beam pickup signal
dpipe: 25 mm
dbutton: 10 mm
400 MHz delay-line filter network
conditioned analog signal
25 ns
25 ns
up
up
down
down
12.5 ns delay-line

## Slide 81: Waveform Sampling Options
4 GSPS = 250 ps sample spacing
  100 samples per 25 ns with single ADC
    50 samples per electrode signal, but only 20…24 samples ≠ 0
  200 samples per 25 ns with two interleaved ADC in I/Q operation
Requires low CLK jitter within 25 ns time interval
  E.g. ta = 0.15 ps -> 68.5 dB@400 MHz (equivalent to EOB=11.1)

## Slide 82
Schematic of EO-BPM
See also: S.Gibson (RHUL this conference)

## Slide 83
EO-Crystal
Analyser
(in-fibre) Polariser
laser light
to fibre coupling
GRIN lenses
All optical BPM laser lab set-up

## Slide 84
Wide-Band Improvement on RF Hybrid Junction: (MSM)
Sum Σ and Difference Δ signals are computed in electro-optical domain
  Aim at 12+ GHz Bandwidth
Possible Detection scheme

## Slide 85: Instead of a summary: Thank You for Your Patience