## Slide 1: Development of the beam loss monitor for beam halo measurement in the J-PARC RCS
Masahiro Yoshimoto* , Hiroyuki Harada, Shinidhi Kato, Kota Okabe, Michikazu Kinsho Japan Atomic Energy Agency, JAEA / J-PARC

## Slide 2: Outline
Introduction
Transverse beam halo measurement in the RCS.
  Wire scanner type with beam loss monitors
  Correction formula
  Reconstruction of the beam profile
New BLM development for the Intra-bunch beam halo measurement
  Scintillation-type BLM of lead glass
  Cherenkov-type BLM of quartz or UV acrylic
  Trial beam test
Conclusion

## Slide 3: Overview of J-PARC 3GeV RCS
RCS beam parameter
Circumference 348.333m
Superperiod 3
Harmonic number 2
Injection energy 400MeV
Extraction energy 3GeV
Repetition rate 25Hz
Particles per pulse 8.3e13
Average current 333mA
Output beam power 1MW
Injection
section
Extraction
section
RF section
1st arc
section
2nd arc
section
3rd arc
section
from LINAC
to MLF
to MR
400MeV Linac
3GeV Rapid Cycling Synchrotron (RCS)
50GeV Proton synchrotron (MR)
Hadron Experimental Hall
Neutrino Experimental Facility
Materials and Life Science Facility (MLF)
Transmutation Experimental Facility
J-PARC = Japan Proton Accelerator Research Complex
Three accelerators:
  400MeV LINAC / 3GeV RCS / 50GeV MR
Multi purpose research facility with high power proton beam
  Nuclear and Particle Physics
  Material and Life Science
Configuration of the Accelerator Complex of J-PARC
3GeV - RCS
multi-intermediate pulse structure formed by the RF chopper
2 bunched structure
Extraction scheme:
  Fast extraction with kicker and septum magnet
H-
H+
Injection scheme:
  Charge-exchange multi-turn injection with stripping foil

## Slide 4: Beam Halo Measurement
The J-PARC RCS successfully achieved 1-MW beam acceleration in January 2015.
To provide such a high power proton beam for the MR and MLF, it is required to improve the extraction beam quality, namely to achieve the Low-Halo and High-Intensity beam by finer beam tuning in the RCS.
Therefore the measurement of the transverse beam profile including both of the beam core and the beam halo is one of the key issues for the high power beam operation in the RCS.
Beam halo monitor
Beam halo monitor
RCS beam halo monitor is installed at the extraction beam transport line (3NBT).

## Slide 5: Beam Halo Monitor in the 3GeV-RCS
Beam halo monitor is combined a wire scanner and several beam loss monitors.
By using several BLMs with the various sensitivities, the ultra-wide dynamic range can be achieved.
Several beam loss monitors
with various sensitivities
wire scanner scraper
radiation
Extraction
beam
Vertical
scanning device
Horizontal
scanning device
BLM detects the whole radiations which are generated at the various place around the monitor.
By subtracting the background signal from the whole signal, pure beam loss signal at the wire scanner can be obtained.
inserted wire scanner
= whole signal
retracted wire scanner
= background signal
pure beam loss signal at the wire scanner
subtraction

## Slide 6: Various sensitivity of BLMs (I)
Plastic scintillator and photomultiplier (PMT) assemblages were adopted as the BLMs with quick time response.
In order to achieve the ultra-wide dynamic range, we need to control the sensitivities of the BLMs.
By changing the volume of the plastic scintillator, the sensitivity-difference can be achieved.
H=300mm
φ=60mm
L=100mm
W=20mm
Scintillator
PMT: H6410
Large scintillator BLM
Light guide
scintillator
Viewing from interior of the light guide
Scintillator
W=30mm
H=32mm(2mm inside in the light guide）
T=2mm
Small scintillator BLM

## Slide 7: Various sensitivity of BLMs (II)
Another method to achieve the sensitivity difference is changing the distance from the wire scanner to the BLMs.
To MLF TGT
To 3MBT-BD
From RCS
Beam halo monitor
Large PS-BLM
Far from WS
Small PS-BLM
Large PS-BLM
Near from WS
High sensitivity
Low sensitivity
Low sensitivity

## Slide 8: Calibration of Beam Loss Monitors
Injection bump
I-BPM
Injection bump
I-BPM
In this system, BLMs can detect the only radiations caused by hitting the beam into the wire scanner.
Thus BLMs signal should be proportional to the number of the hitting particles.
Calibration formula (Nparticle / SBLM) for each BLM can be obtained by using the matched beam intensity with each sensitivity.
High sensitivity detector:
Low sensitivity detector:

## Slide 9: Reconstruction of the transverse profile
Raw data
After calibration
Using each sensitivity calibration formula
Only by using sensitivity calibration formula for the each BLM, transverse profile including both the beam core and beam halo can be reconstructed directly.
( Any other correction is not necessary to join together all measured plots.)
In this scheme, it is possible to expand the dynamic range by using the higher sensitivity BLMs.

## Slide 10: Intra-bunch beam halo measurement
The beam halo monitor, which is combined with the wire scanner and several BLMs with various sensitivities, can reconstruct the transverse beam profile including both the beam core and halo.
Our final goal of this halo monitor is also to measure the intra-bunch beam halo of extracted two bunches from the RCS.
Wire scanning
Transverse beam profile
Intra-bunch beam profile
time
BPM signal
projection
projection

## Slide 11: Reconstruction 2D-profile
2D-profiles ( bunch length vs beam size ) were reconstructed from the two BLMs with low sensitivity.
Two 2D-profiles differ in the time distribution. ⇒ why ?
Near Small
PS-BLM
Far Large
PS-BLM

## Slide 12: Distortion of the bunch signal
The plastic scintillator BLMs detected not only the radiation from the wire but also reflected one from other devices and wall.
Wire scanner
MLF-TGT
3GeV-RCS
( 3GeV proton beam )
To obtain the bunched beam structure from the BLM clearly, it is necessary to suppress these reflected radiation.
Therefore we develop new-type BLMs, which are scintillation-type BLM of lead glass and Cherenkov-type BLM of quartz or UV acrylic.
subtraction
subtraction

## Slide 13: new-type BLMs
Scintillation-type BLM of lead glass
These new-type of BLMs have high sensitivity toward the faster gamma-ray or electron.
Thus the suppression of the reflected radiations may be expected.
30mm
f40mm
60mm
f40mm
QG
UVA
QG
UVA
QG
UVA
Cherenkov-type BLM of quartz or UV acrylic
200mm
20mm
19mm
LG
LG
LG
PMT: H6410
PMT: H6521
PMT: H6521

## Slide 14: Trial beam test with new-type BLM
To see if the new-type BLMs work as they were expected to, they were installed near the wire scanner and trial beam tests were carried out.
(Particle simulation with GEANT4 is starting now.)
Wire scanner
MLF-TGT
3GeV-RCS
( 3GeV proton beam )
To enhance the effect of the reflected radiation, the beam destination from the RCS switched from MLF-Target to 3NBT-Beam Dump.
At first, the ultra-low intensity beam was used to prevent the saturation of the BLM.
Next, the high intensity beam was used for the high power beam commissioning.
3NBT-BD
QG
UVA
LG

## Slide 15: Ultra-low intensity beam test
Ultra-low intensity beam (1-intermediate bunch with thin chopped width)
By switching the beam destination from MLF-TGT to 3NBT-BD
All new-type BLMs can suppress the reflected radiation drastically.
Plastic scintillator
UV acrylic
Quartz Glass
Lead Glass

## Slide 16: High intensity beam test (I)
High intensity beam test (747kW-eq.)
Bema destination is fixed to 3NBT-BD, compare between inserted and retracted the wire scanner.
All bunch signals measured by new-type BLM were almost not distorted.
⇒ The transverse profile can be measured in the 3NBT-BD operation.
Plastic scintillator
Lead Glass
UV acrylic
Quartz Glass

## Slide 17: High intensity beam test (II)
Transverse beam profiles were measured with new-type BLM in the 3NBT-BD operation with various beam intensity (138 ~ 747kW-eq).
The Plastic scintillator signal was saturated due to the reflected radiation from the 3NBT-BD, and transverse profile was distorted.
New BLM signals can obtained the beam profiled.
Plastic scintillator
UV acrylic
Quartz Glass
Lead Glass

## Slide 18: Trial of the reconstruction
Calibration formula of the each new-type BLM were calculated roughly
The transverse beam profiles were reconstructed .
Transverse beam profile can be obtained in the 3NBT-BD operation.
⇒ It is good advantage for beam commissioning in the RCS.
Rough calibration formulas
Beam power: 747[kW-eq.]
550[kW-eq.]
343[kW-eq.]
138[kW-eq.]

## Slide 19: conclusion
The transverse beam profile including both the beam core and beam halo can be reconstructed by using the several BLMs with the various sensitivities.
Only by using sensitivity calibration formula for the each BLM, transverse profile including both the beam core and beam halo can be reconstructed directly.
In this scheme, it is possible to expand the dynamic range by using the higher sensitivity BLMs.
The time structure of the bunch signal measured by the plastic scintillator was distorted, because of the reflected the radiation.
The new-type BLMs, which are scintillation-type BLM of lead glass and Cherenkov-type BLM of quartz or UV acrylic, can be reduced the reflected radiation drastically.
In order to achieve the intra-bunch beam halo measurement, the new-type BLMs have to progress.
Thank you for your attention !