## Slide 1: SNS Beam Diagnostics: Ten Years After Commissioning
A. Aleksandrov
Oak Ridge National Laboratory,
USA

## Slide 2: Spallation Neutron Source Accelerator
Front-End:
Produce a 1-msec long, chopped, H- beam
1 GeV LINAC
Accumulator Ring: compress 1-msec long pulse to 700 nsec
2.5 MeV
Warm LINAC
Front-End
RTBT
HEBT
Injection
Extraction
RF
Collimators
945 ns
1 ms macropulse
Current
mini-pulse
Chopper system makes gaps
Current
1ms
Liquid Hg Target
1000 MeV
1 ms macropulse
1 ms
<1 msec
Cold LINAC
186 MeV
38mA
38A

## Slide 3: History plot of SNS beam power
design power
BIW2010, SNS Beam Diagnostics Experience and Lessons Learned
~ 4700 hours/year

## Slide 4: Second Target Station Project: Double beam power to 2.8 MW
Increase energy from 1GeV to 1.3GeV by adding 36 Super Conducting linac cavities
Increase beam current by 50%
Enables second target station in 2020s
Beam Instrumentation Group Involvement
  more of the same: BCM, BPM, WS
  Intra-bunch ring feedback system
  laser stripping injection

## Slide 5
SNS Beam Instrumentation Systems
are Numerous, Diverse and Growing in Number
RING
44 Position
54+ Loss
1 Current 12 Fast Loss
5 Electron Detectors
2 Electron Profile Scanner
2 Transverse BTF
SCL
34 Position and Phase
33+ Loss
9 Laser Wire
24 Neutron Detectors
HEBT
29 Position and Phase
26+ Loss,
3 Fast Loss
10 Wires
4 Current
2 Bunch Shape
6 Scrapers Charge
1 Laser emittance
IDump
3 Position
1 Wire 6 Loss
2 Current
2 Video
EDump
1 Current
4 Loss
1 Wire
LDump
6 Loss
6 Position
2 Wire
CCL
10 Position and Phase
9 Wires
48 Loss
1 Faraday Cup
4 Bunch Shape
10 Neutron Detectors
MEBT
6 Position and Phase
2 Current
5 Wires
1 CHUMPS
1 Emittance
1 Laser longitudinal profile
DTL
10 Position and Phase
5 Wire
4 Loss
5 Faraday Cup
6 Current
18 Neutron Detectors
RTBT
17 Position
26 Loss
5 Wires
4 Current
1 Harp
3 Fast Loss
1 Target Imaging
MDump
4 Loss
2 Current
1 Wire
BCM, BLM, BPM, BSM, WS….
15+ systems, 400+ IOCs

## Slide 6: SNS Beam Diagnostics Performance Assessment
 | Function | Neutron Production | Machine Tuning | Single particle model | RMS model | Hi Res model | Obsolescence problem
BCM (18) | current |  |  |  |  |  | 
BLM (360) | radiation |  |  |  |  |  | ♪ 0.5M$
BPM (160) | Position x, y, z |  |  |  |  |  | ♪ 1.5M$
Target Harp (1) | Transverse profile, 1D |  |  |  |  |  | 
WS (40) | Transverse profile, 1D |  |  |  |  | ♪ | 
Laser WS (10) | Transverse profile, 1D |  |  |  |  |  | 
BSM (6) | Longitudinal profile, 1D |  |  |  |  |  | 
Laser BSM (1) | Longitudinal profile, 1D |  |  |  |  |  | 
2.5 MeV emm (1) | Transverse emittance, 2D |  |  |  |  |  | 
1 GeV emm (1) | Transverse emittance, 2D |  |  |  |  | ♪ | 
Model based machine tuning
now
1-2 years
3-5 years

## Slide 7: Beam Position and Phase Monitors (BPMs)
Main tool for machine tuning and troubleshooting
  Phase measurements for linac tune-up
  Position measurements for trajectory correction, injection set-up and centering beam on dumps and target
160 strip-line pick-ups
  96 “linac type” operate at 402.5MHz and 805MHz
  64 “ring type” operate at low frequency, 5MHz BW
Custom made PCI analog front-end and digital cards
LabView software under embedded Windows XP on individual PCs (one per pick-up), 1Hz trigger rate
Meets all accuracy specs but reliability is not stellar
Hardware obsolescence is major problem
  Parts, cards, PC motherboards, OS upgrades
Short term solution: stock up on spares
Long term solution: new system
PCI board
BPM PC chassis
BPM pick up

## Slide 8: Our approach to all new electronics designs is to minimize in-house design efforts
Analog Front-End
Analog-To-Digital
Conversion
General Purpose FPGA
Crate and controller
In-house design
Commercial Off-The-Shelf (COTS)
High Level Language Programming

## Slide 9: Linac BPM electronics in the field
Old BPM electronics
1 BPM per chassis
New BPM electronics
6 BPMs per crate
AFE
Digital
AFE and digital

## Slide 10: New BPM electronics tested with beam
Courtesy of C.Long and A.Shishlo

## Slide 11: Ring BPM prototype electronics
8 BPMs per crate
Parallel “high” – “low” gain channels combined in FPGA to cover 60dB dynamic range (instead of fast gain switching in the current design)
AFE prototype developed
One BPM set tested with beam
PCB is being designed
Courtesy of R. Dickson

## Slide 12: Beam Loss Monitors (BLMs)
Major tool for machine protection and tuning
Ionization Chamber Detectors (307)
Scintillation Detectors (55)
  Neutron detectors
  Fast loss detectors
Multi-channel analog front-end VME cards
Digital electronics in VME crate
VxWorks software
Very reliable
Hardware obsolescence is becoming a problem
Short term solution: stock up on spares
Long term solution: new electronics

## Slide 13: “Next BLM” system concept
Detectors and cabling shell stay the same
Analog electronics should be as simple as possible
  Four different flavors of chassis
  Two flavors of custom 4-channel front end card
Unification of the chassis, boards, power supplies
Flavor | Purpose | # Signals | # HV | # MPS | Amplifier | Comment
ITSF | PMTs in ITSF | 8 | 8 | 1 | None | Will have just one MPS channel
Ion Chamber | Regular BLM | 16 | 4 | 16 | IC Amp | Standard IC in accelerator
Target BLM | Target Facility | 8 | 8 | 0 | Target Amp | Sensitive DC amplifier for target people
Neutron Detector | NDs | 8 | 8 | 8 | IC Amp | Standard ND in accelerator
Amplifier | Gain (Ohm) | Min Current (nA) | BW (Hz) | Sampling (kS/s)
IC Amp | 600k | 2 | 200k | 1000
Target Amp | 100M | 0.01 | 1 | 100
Courtesy of A. Zhukov
Output A
Output B
Output C
Output D
Sig A
Sig B
Sig C
Sig D
Power

## Slide 14: 4U Chassis layout for ITSF
Neutron Detector
Ion Chamber
ITSF PMT
Courtesy of A. Zhukov

## Slide 15: 8 channels PMT BLM chassis in the field

## Slide 16: High resolution PIC model development
Request for emittance measurements (2-D at least)
  At as many locations as possible: 2.5MeV, 1GeV, in between
  High dynamic range: 104 - 106
2.5MeV measurement
with scraper out
2.5MeV measurement
with scraper in
1 GeV measurement (sort of) with scraper out (red) and in (blue)
An example of beam transport measurement with ~103 dynamic range

## Slide 17: Available large dynamic range diagnostics
2.5 MeV slit-slit emittance scan
  104-105 dynamic range
  20ns temporal resolution
1 GeV laser emittance scan
  103 dynamic range
  10ns temporal resolution
Wire scanners
  105 dynamic range
  50us time resolution
Need to learn using wire scanners for phase space measurements

## Slide 18: MENT Tomographic Reconstruction of 2-D Emittance from 1-D Profiles
Reconstruction seems to works very well in HEBT
  Need to verify using laser emittance measurements
  High resolution reconstruction requires iterative procedure. 103 dynamic range demonstrated
Plan to extend to SCL, Warm Linac, MEBT
  Requires good transport model
  Problem of space charge
Comparison of measured and reconstructed profiles
Reconstructed 2-d distribution

## Slide 19
Many thanks to SNS Beam Instrumentation Team members who provided material for this talk:
Wim Blokland, Richard Dickson, Cary Long,
Yun Liu, and Sasha Zhukov
and
Thank you for your attention!