## Slide 1: Fast-Scanning Ptychography
and other scanning probe methods
Cameron M. Kewish
cameronk@ansto.gov.au

## Slide 2: Outline of the presentation
Introduction to ptychography and recent developments
Fast-scanning ptychography at the XFM Beamline
  Synchronising data streams
  User interface
  Behind the scenes
  Fast scanning achievements
Application examples: geophysical exploration, fatigue cracking
New mapping capabilities: micro-beam SAXS & diffraction
Capabilities of the BRIGHT Nanoprobe

## Slide 3: Introduction to ptychography
Super-resolution microscopy by moving aperture
P. D. Nellist, et al., Nature 374, 630 (1995)
J. M. Rodenburg & R. H. T. Bates, Phil. Trans. R. Soc. Lond. A 339, 521 (1992)
H. M. L. Faulkner & J. M. Rodenburg, PRL 93, 023903 (2004)

## Slide 4: Introduction to ptychography
Crystals illuminated with tightly focused beams produced overlapping Bragg peaks, which allowed analytical phasing
The data include both aperture and object functions ‘folded’ together
W. Hoppe & G. Strube, Acta. Cryst. A 25, 502 (1969)
R. Hegerl & W. Hoppe, Ber. Bunsen-Ges. Phys. Chem 74, 1148 (1970)
P. D. Nellist, et al., Nature 374, 630 (1995)
W. Hoppe, Acta. Cryst. A 25, 508 (1969)

## Slide 5: Introduction to ptychography
SXDM, or ptychography
fluorescence photons
sample scanning stage
X-ray detector
monochromatic X-rays
focusing lens
photo-
electrons
STXM, SPEM, XFM
CCD camera
X-rays
condenser zone plate
TXM
coherent X-rays
focusing optics
CCD or hybrid pixel
array
coherent diffraction
fluorescence photons
objective
zone plate

## Slide 6: Imaging at hierarchical length scales
electron microscopy
visible light microscopy
scanning probe microscopy
X-ray radiography & tomography
X-ray crystallography
X-ray microscopy & µCT
X-ray coherent diffractive imaging
Synchrotron
Laboratory
Hospital
(Hard) X-ray advantages:
High penetration power
Chemical specificity
No vacuum required
1 Å 10 nm 1 µm 1 mm 1 m

## Slide 7: Developments in the last decade
Review: F. Pfeiffer, Nat. Photonics 12, 9 (2018)
Ptychographic X-ray computed tomography
Bragg ptychography
Fresnel (near field) ptychography
Spectroptychography

## Slide 8: Developments in the last decade
Ptychographic methodology developments: PXCT, Fresnel, Bragg, …
Algorithmic developments
  Difference Map with probe retrieval: Thibault (2009)
    optics characterisation & alignment: Kewish/Guizar-Sicairos/Schropp (2010)
  Nonlinear Optimisation:
    maximum likelihood refinement: Guizar-Sicairos (2011)
  “Ptychographical Iterative Engine” (PIE):
    PIE variants: pPIE, ePIE, mPIE, 3PIE, etc.: Rodenburg group/PhaseFocus
  Mixed states solutions: Thibault & Menzel (2013)
Experimental advances:
  position refinement, resonance, multimodal, flyscans…

## Slide 9: Optimised bidirectional scanning
Raster step scan
Unidirectional or bidirectional
Movement overhead ~0.1-1s/point
Up to 99% of scan time wasted
Large area scans unfeasible
Line flyscan
Fast flyback or step to next line
Movement overhead ~0.3-2s/line
Up to 6-30% of scan time wasted
Rascan
Mathematically optimized trajectory
Turnaround time ~0.045s/line
Less than 1% of scan time wasted
Megapixels per hour is feasible
N. Afshar et al., Proc. ICALEPCS2017, WEBPL07 (2017)
Example: 50 x 50 µm², 100 nm step, 10 ms dwell
Total measurement time: 500² x 0.01 s = 2500 s

## Slide 10
XFM: X-ray Fluorescence Microscopy
Fast-scanning hard X-ray microprobe – mapping chemistry:
MAIA detector in backscatter geometry
4-27 keV energy range (up to Cd-K)
Milliprobe: Secondary source aperture defines 100 µm beam
Microprobe: KB-mirrors focus to approx. 2×2 µm beam size
Flyscanning megapixels per hour, XANES imaging
Double Multilayer
Monochromator
(coming soon)
See recent XFM beamline article in JSR:
“The XFM Beamline at the Australian Synchrotron”
D. L. Howard, et al., J Synchrotron Rad 27, 1447 (2020)
Shielded transport
“Milliprobe”
instrument
Secondary
Source
Aperture
Microprobe
instrument
Small KB Mirror System
(in endstation)
recently refurbished
with new mirrors, motors
and enclosure

## Slide 11
Incident beam
MAIA 384D detector
Transmitted beam continues
to SXDM
detector
High speed servomotor
sample scanning system
VORTEX ME3
detector
K-B mirrors
XFM: Microprobe instrument schematic
Beam is focused by
K-B mirrors onto the sample, which is scanned through the micron-sized beam while detectors record fluorescence and diffraction

## Slide 12: Ptychography installation at XFM
SXDM detector positioned as far as achievable (3.85 m) downstream of sample
X-ray beam travels inside a helium-filled ‘flight tube’
Sample is scanned orthogonal to beam behind the MAIA detector permitting simultaneous XFM + SXDM experiments
Detector events are recorded into the MAIA data stream
He tube
SXDM detector
sample
coherent
focused
X-ray beam

## Slide 13
Cr XFM
Ptychography absorption contrast
2015: Dectris loaned a EIGER 1M detector
Flyscan ptychography data was measured
simultaneously with XFM data
Correlated XFM:Ptychography test
M. W. M. Jones et al., J. Synchrotron Rad. 23, 1151 (2016)
Imaging
length
scale:
bars are
0.5 µm
wide at
inner
edge of
“0.5” ring

## Slide 14: Flyscan at Australian Synchrotron: XFM
 | STEP SCAN Slow point scan
Sample motion | Stop-motion
EIGER2 Trigger mode | INTS
Trigger behaviour | 1 image per trigger, N triggers in scan
Trigger type | Software trigger from EPICS scan record
Position data | List of N positions
Position uncertainty | =20 nm, =20 nm
Practical dead time | 0.3 s per pixel
 | STEP SCAN Slow point scan
Sample motion | Stop-motion
EIGER2 Trigger mode | INTS
Trigger behaviour | 1 image per trigger, N triggers in scan
Trigger type | Software trigger from EPICS scan record
Position data | List of N positions
Position uncertainty | 
Practical dead time | 0.3 s per pixel
FREE RUN Camera driven data
Continuous
EXTS
N images per trigger, first starts the scan
TTL first pixel, then internal clock
Encoders recorded into event stream
=10 nm, =10 nm
Practically zero
FREE RUN Camera driven data
Continuous
EXTS
N images per trigger, first starts the scan
TTL first pixel, then internal clock
Encoders recorded into event stream
Practically zero
RASCAN “Standard” XFM
Continuous
EXTS
1 image per trigger, N triggers in scan
TTL from HYMOD pixel entry events
Pixel triggers recorded into event stream
=20 nm, =½ pixel
45 ms/line +10% dwell*
RASCAN “Standard” XFM
Continuous
EXTS
1 image per trigger, N triggers in scan
TTL from HYMOD pixel entry events
Pixel triggers recorded into event stream
45 ms/line +10% dwell*

## Slide 15: List mode: synchronising data streams
X Pixel transit event TTL
  EIGER “EXT IN” = start acquisition
EIGER “EN OUT” = acquisition in progress
  Rising edge = frame event
  …
  End of scan
“Frame events” are recorded in list mode with MAIA data: provides 1:1 correspondence between EIGER frames and scan pixels.
New May 2019: “Stage encoder” events implemented by CSIRO for subpixel synchronisation of positions
pixel array
20 nm encoder grid
Rascan trajectory
X
X
X
X
X
X
X
X
X
X
X
N. Afshar et al., Proc. ICALEPCS2017, WEBPL07 (2017)
C. G. Ryan et al., Synchrotron Radiation News 31, 21 (2018)
“Pixel Triggered” acquisition

## Slide 16: Camera-driven acquisition:
X Trigger sent to EIGER at scan start
  EIGER “EXT IN” = start acquisition
EIGER “EN OUT” = acquisition in progress
  Rising edges = frame events
  …
  End of scan
Avoids race condition of arriving at the next pixel while still acquiring, due to velocity variations
Position symmetry is broken, avoiding “raster grid pathology” artefacts in ptychography
Freedom to alter frame length, changing overlap between X position compared to pixel array
pixel array
20 nm encoder grid
Rascan trajectory
X
“Freerun” acquisition

## Slide 17: Position Interpolation
Pixel
Encoder

## Slide 18: Pixel triggering vs freerun acquisition
Exposure time and pixel transit time creates race condition:
Small variations in scan velocity can lead to missed triggers
pixel triggered scan: frame_time set to 0.99 × dwell
freerun scan: frame_time set to 1.00 × dwell

## Slide 19: Correlated XFM + SXDM
C. M. Kewish et al., SPIE Nanophotonics 10456, 51 (2017)
SXDM implemented at XFM with an X-ray photon counting detector in 2017
Mouse lung: 10 µm-thick section provided by Godfrey group, Doherty Institute
250 μm
XFM – K
100 nm pixels
2 μm
SXDM – phase
25 nm pixels
XFM – K
10 µm pixel size

## Slide 20: Upgrade detector – EIGER X at XFM
Original detector was cost-effective but slow, data poor, analysis troublesome
  Slow readout: significant dead-time between exposures
  Slow dynamics: XFM beam had to be heavily attenuated, wasting beam time
  Sensor charging: intensity- and time-varying bias, frequent power cycling
Collaboration: ANSTO – La Trobe University
  Provide an EIGER2 X 1M detector to the XFM beamline
    Loan EIGER1 detector provided by Dectris during 2018 during development of EIGER2 X
  Develop an optimised automated data analysis pipeline
    Roll-out of software in 5 Phases, providing prompt access to SXDM for XFM ‘early adopter’ users
  Provide computing infrastructure for SXDM
    Collaboration with IT/SciComp to optimise hardware/software for HPC imaging

## Slide 21: Upgrade to EIGER2
 | Original detector | EIGER1 X | EIGER2 X
Date | Delivered 2016 | Loan June 2015 (4 weeks) Loan May 2018 (48 weeks) | Delivered May 2019
Area [mm2] | 31 × 25 | 77.2 × 79.9 | 
Readout time [µs] | 8,000 | 3 | 0.1
Max intensity per pixel [photons/s] | 105 | 106 | 107
Robustness | Poor: Power cycling to disperse charge every ~1 h takes ~5 min, crashes if triggered during acquisition, cannot transfer data during scan | Excellent: Acquisitions >100 million frames: OK Stable operations >8 weeks without initialization | 
Capabilities enabled | SXDM, STXM | SXDM, STXM, µ-diffraction, µ-SAXS | 

## Slide 22: Upgrade – Data Analysis Pipeline
The major steps in SXDM analysis comprise:
confirm the X-ray beam intensity and illumination conditions pre-scan;
carry out one or more scans using XFM maia_scan IDL architecture;
transfer EIGER data from detector control unit (DCU) to beamline fileserver;
extract pixel visit data and/or stage events from the list-mode data;
convert positions to laboratory coordinates using metadata from maia_scan;
locate the beam centre in EIGER frames;
apply valid pixel mask, and extract region of interest (ROI) from each frame;
save list of positions and EIGER ROIs as input for SXDM reconstruction, and;
reconstruct SXDM images.

## Slide 23: Upgrade – Integration & Automation
What users & scientists see:
EPICS areaDetector control GUI
EIGER realtime view
EIGER Data Analysis GUI
Confluence pages
Most of the program code is hidden to flatten the learning curve for the users
XFM Users configure scans as normal in Excel sheet, maia_scan with “/camera” keyword to activate EIGER

## Slide 24: Upgrade – Integration & Automation
Behind the scenes:
EIGER temp & humidity
status logging
live mode status
live view console
data transfer daemon
status logging
EPS status IOC
live mode control IOC
ADEiger MEDM console
areaDetector:ADEiger IOC
xfmpro01
xfmpro02
cluster status
xfmpro03
Control and monitoring software runs on a virtual host
Data analysis pipeline is implemented both on ASCI and the SXDM cluster “xfmpro”
ASCI is accessible from anywhere
xfmpro is currently accessible only from onsite/VPN

## Slide 25: Scanning transmission X-ray microscopy
STXM - resolution (point spread function) defined by beam
  image size determined by scan points
  each frame yields 1 pixel:
Transmission
  integrated intensity in each frame
Darkfield
  integrated scatter outside beam
Differential phase contrast (DPC)
  vertical and horizontal displacement
  gradient of the refractive index
transmission
darkfield
horizontal DPC
vertical DPC
A. Menzel et al., Ultramicroscopy 110, 1143 (2010)

## Slide 26: Fast scanning achievements
Chart: Scan area coverage [mm²/min]
| EIGER2 - Test pattern 0.4 ms -- 200 nm resolution | EIGER2 - Geosample 3 ms - <100 nm resolution - 2.6 x 1.3 mm² | EIGER2 - Test pattern 6.25 ms - <50 nm resolution | EIGER1 - Nematode worm 20 ms - 280 nm resolution - 0.8 x 0.09 mm² | EIGER1 - Geopolymer 25 ms - 65 nm resolution | EIGER1 - Microfabricated sample 50 ms - 70 nm resolution | Previous detector - Test pattern 100 ms - 200 nm resolution
area/time | 0.05 | 0.015146 | 0.0035573 | 0.00044199 | 8.1818e-05 | 1.1667e-05 | 8.3333e-07
Scanning speeds (and data quantities) have increased by orders of magnitude
In development, 20 x 20 µm² area could take >1 hour, at 1 µm/s scan velocity
EIGER2 high intensity capability and dead-time free framing permits >250 µm/s
Gigapixel nanoscale imaging of geosamples and biosamples is achievable
10 Hz
20 Hz
40 Hz
50 Hz
160 Hz
300 Hz
2500 Hz
Frame rate
5 gigapixel ptychography
image acquired in 179 min

## Slide 27: Large area, high resolution
20 µm
SXDM phase – 25 nm pixels
100 µm
19 micron thin section of geosample
Scan area: 800 × 440 µm²
Scan velocity: 150 µm/s
Exposure time: 3.33 ms (300 Hz)
Total scan time: 79 min
Total data: 1.4 million diffraction patterns
Ptychography total image size: 523 megapixels
XFM - FeMnK
STXM darkfield – 0.5 µm pixels
Sample from C. Schrank (QUT)

## Slide 28: Fast scanning ptychography
STXM transmission
Each EIGER frame is integrated to give a transmission pixel.
Average intensity: 14500 photons/frame including direct beam
8-bit readout:
intensity was
<6e5 ph/pixel/sec
100 µm
300 nm-thick Au-Cr Siemens star
Scan area: 300 × 350 µm²
Scan velocity: 250 µm/s
Exposure time: 0.4 ms (2500 Hz)
Total scan time: 12.5 min
Total data: 1.86 million diffraction patterns
SXDM phase
0.2 µm ring can be clearly resolved
from scans with 0.4 ms exposures
All features including <50 nm are resolved with 20 ms exposures
14 nm pixels
SEM

## Slide 29: Geoscience – Correlated XFM & phase
Formation of calcite veins in shale:
Solid arrows in C and D show a break in the nano-crystallites accommodating the fibrous calcite growth that coincides with small Zn-rich grains.
Thin arrows highlight an interface of calcite grains with euhedral quartz crystals decorating the vein surface. This interface has a width of ~250 nm and served as fluid pathway for Cu mineralization.
M. W. M. Jones1, C. Schrank1, G. A. van Riessen2, C. M. Kewish3
1 Queensland University of Technology
2 La Trobe University
3 ANSTO Australian Synchrotron
SXDM
47 nm pixels
XFM
100 nm pixels
XFM – Ca
10 µm pixels
microscope

## Slide 30: Dynamic processes – fatigue cracking
Fatigue fracture in sub-micron Ni thin films:
Ni film on thin Kapton, secured in a loading rig. One end is attached to a stepper motor, controlled via EPICS IOC on a Raspberry Pi.
SXDM repeated after N load-unload cycles consisting of X% film length increase. Time between scans is 3 minutes.
Darkfield quantifies total scatter at each point
A sub-micron crack is observed appearing
G. N. Hinsley1, G. A. van Riessen1, C. M. Kewish2
1La Trobe University
2ANSTO Australian Synchrotron
(manuscript in preparation)
SXDM – phase
50 nm pixels
crack
STXM – darkfield
300 nm pixels

## Slide 31: New XFM capabilities
Preliminary result: pressure dependence in peaks from monoolein (LCP) Pn3m phase
HPLC pressure data recorded with each EIGER2 frame
In addition to SXDM, EIGER2 enables: µ-diffraction and µ-SAXS experiments
Detector distances from 3.88 m (SXDM/µ-SAXS)  0.03 m (<3Å µ-XRD) have been used
EIGER2 data
online view
sample
injector
µSAXS
profile
P. Berntsen1, C. Darmanin1, B. Abbey1, T. Greaves2, A. Martin2, C. M. Kewish3
EIGER2
sample
0.8 m
HPLC pump
Background
corrected
EIGER2
µSAXS
data
XFM setup
XFM controls
stable sample jet with
ASU
injector
1 LTU
2 RMIT
3 ANSTO

## Slide 32: ANSTO Australian Synchrotron
10 Beamlines: Operational since 2008
IMBL: Imaging and Medical Therapy
IRM: High-Resolution Infrared Microspectroscopy
THz: Terahertz/Far-Infrared Fourier Transform Spectroscopy
PD: Powder Diffraction
MX-1/2: Macromolecular Crystallography 1 & 2
SXR/SXR-I: Soft X-ray Spectroscopy & Imaging
SAXS-WAXS: Small Angle X-ray Scattering/Wide Angle X-ray Scattering
XAS: X-ray Absorption Spectroscopy
XFM: X-ray Fluorescence Microspectroscopy & Ptychography
3 GeV machine: 4000+ research visits/year
Project BRIGHT: 2018-2026
MCT: Micro Computed Tomography
MEX-1/2: Medium Energy X-ray Absorption Spectroscopy 1 & 2
BioSAXS: Biological Small-Angle X-ray Scattering
ADS-1/2: Advanced Diffraction and Scattering 1 & 2
MX-3: Macromolecular Crystallography 3
Nanoprobe: Multimodal Fluorescence Nanospectroscopy & Ptychography
https://www.ansto.gov.au/research/facilities/australian-synchrotron/project-bright
Nanoprobe Satellite Building
ADS Building

## Slide 33: Nanoprobe Capability targets
Diffraction, nano-WAXS
Industrial insights
Beam stop
Window scatter
(isotropic)
WAXS of (0002) graphitic planes aligned along the fibre
SAXS of ellipsoidal microvoids
  Multi-element mapping @ sub-100-nm resolution
  XANES fingerprinting of chemical coordination environment
Zn K Se
Fluorescence tomography
Ultra-high resolution ptychographic imaging
  Elemental Mapping
  Advanced high-resolution methods
  XANES spectroscopy
  Complementary & emerging methods
  Elemental Tomography

## Slide 34: Mapping sensitivity & resolution
Native contrast: no dyes or contrast agents necessary
can be used to provide functional/biomarker information
Quantitative
Non-destructive / minor damage
Extended penetration & depth-of-field:
intact cells, organisms & tissue sections
Simultaneous mapping of 10+ elements:
Hard X-ray range: 5-17 keV (DMM)
K-edge: Si – Sr; L-edge: ~Ag – Po
Extremely High sensitivity: ppb; sub-µM
Sensitive to chemical speciation via XANES (DCM)
K-edge: Ti – Sr; L-edge: Cd(LIII) – Po
LA-ICP-MS
Nanoprobe
Sensitivity
ppb
ppm
0.1 μm
1 μm
10 μm
100 μm
LMD-LA-ICP-MS
SEM-EDX
PIXE
EJ New Dalton Trans (2013), 42(9) pp 3210
Spatial resolution
XFM (today)
XFM (DMM)
Notes: This sensitivity corresponds to about 120,000 Fe atoms within the ~200 nm focus

## Slide 35
Quantitative 3D elemental microtomography of Cyclotella meneghiniana at 400-nm resolution
PNAS (2010)
35 h of measurement in 2010
Targeting 8 h at Nanoprobe

## Slide 36: Ptychography – resolution beyond the probe size
Elemental distributions
  90 nm
Chlamydomonas (green algae) frozen-hydrated
  ~25 nm
Ptychographic imaging of ultrastructure provides key correlative insights
J. Deng, et al., Scientific Reports 7, 445 (2017)

## Slide 37
  User Communities by Field-of-Research
  Earth Sci
  Bio
  Enviro
  Chem
  Medical
  Physical
  Engineering
  Technology
  Agriculture
  Cultural Heritage
  A large cross-section of XFM users can explore their samples in higher resolution at Nanoprobe
  XFM user FoRs

## Slide 38
ADS
Nanoprobe
IMBL

## Slide 39: Capability targets
Insights into Industrial processing
Beam stop
Window scatter
(isotropic)
WAXS of (0002) graphitic planes aligned along the fibre
SAXS of ellipsoidal microvoids
  Multi-element mapping @ sub-100-nm resolution
Ultra-high resolution ptychographic imaging
  XANES fingerprinting of chemical coordination environment
Zn K Se
Fluorescence tomography
  Elemental Mapping
  Advanced high-resolution methods
  XANES spectroscopy
  Complementary & emerging methods
  Elemental Tomography
  A high-end x-ray fluorescence Nanoprobe will commence operation @ the AS in mid-2024.
  The Nanoprobe will bring multiple tools for interrogating high-nano-scale elemental distributions (Z>14, Si) and ultra high resolution ultrastructural imaging
  Now is the time to develop capabilities & skills, and to establish requirements with the Nanoprobe.
  email: as_nano@ansto.gov.au
  website: https://www.ansto.gov.au/x-ray-fluorescence-nanoprobe-beamline

## Slide 40
Acknowledgments
D. Paterson, D. Howard, N. Afshar, M. de Jonge, J. Reinhardt
ANSTO: Controls, Sci Comp, IT, SAXS, Elec Eng, Stores, PMO, Finance, Legal
CSIRO MAIA team: R. Kirkham, C. Ryan
La Trobe University: van Riessen group, Abbey group
PtyPy Developers: B. Enders, P. Thibault, A. Parsons, et al.
XFM Ptychography Power Users:
G. van Riessen & G. Hinsley, M. Jones, N. Phillips
Early Adopter Users of ptychography at XFM:
La Trobe, QUT, RMIT, Deakin, ANU,
Univ. Paris, Univ. BC
BRIGHT Nanoprobe team
Questions & comments:
cameronk@ansto.gov.au