## Slide 1: Is there potential for a flatter potential?
ATF IC-CoP 17th to 19th October 2022
CAS-ASD
David Button

## Slide 2: What are tandem accelerators and why the question about potential?
Tandem Accelerators are used to accelerator ion beams to a deterministic energy
Ion beams are used for;
  Accelerator Mass Spectrometry (AMS)
  Ion Beam Analyses (IBA)
  Materials Modification
  Irradiation

## Slide 3: Some previous work 
Terminal Voltage Stabilization of Pelletron Tandem Accelerator
N. R. Lobanov, M. C. Blacksell, P. Linardakis, and D. Tsifakis, “Terminal Voltage Stabilization of Pelletron Tandem Accelerator”, in Proc. HIAT'12, Chicago, IL, USA, Jun. 2012, paper TUB03, pp. 124-128.
URL: https://jacow.org/HIAT2012/papers/TUB03.pdf

## Slide 4: Manipulation ions
Electric Field
...............................................................................................................................................
Magnetic Field (Field coming out of the page)
F
v
F
v
F
v
F
v

## Slide 5: So why is stable energy important
Maintaining a known chromatic ion beam energy allows us to achieve;
  Good transmission of the ions from the ion source to the final target
  Accurate filtration and tailoring of the ions
  Deterministic control over the interactions with targets

## Slide 7
Tandem electrostatic accelerator
Tandem accelerators use static electric fields to accelerate particles as typically DC beams.
The Tandem process allows for energies greater than the terminal voltage to be generated.
Ion Source Neg potential deck
++++
Max Voltage
Terminal
Stripper
Ion Source Neg potential deck
Accelerate Negative Ions
q Change
Accelerate Positive Ions
Charging Chains
Charging Chains
SF6 Pressure Vessel
 | Charge State (q) | Acceleration Voltage | Ion Energy Gain
Injection Energy | -1 | 100 kV | 100 keV
Low Energy | -1 | 3 MV | 3 MeV
High Energy | +3 | 3 MV | 9 MeV
Exit Energy |  |  | 12.1 MeV
Low Energy Section
High Energy Section
Efinal = Einj + (1 + q) Vterminal

## Slide 8
Tandem electrostatic accelerator
Ion Source Neg potential deck
++++
Stripper
Ion Source Neg potential deck
Accelerate Negative Ions
q Change
Accelerate Positive Ions
SF6 Pressure Vessel
Low Energy Section
High Energy Section
High Voltage Terminal
Gas Stripper
(Recirculating gas)
Equal Potential Rings
Ion Source
Injection Magnet/Bouncer
Tank (Pressure Vessel)
Analysing Magnet
Switch Magnet
Motor
Generator
Charging Pelletron Chain
Some Parameters
Column Current, LE, and HE
Stripper Pressure, and Gas Type
Up Charge
Down Charge
Inductor and Suppressor Voltage and Current
Loss Charge
Corona Probe Position and Current
Corona Probe Grid Voltage Read Back
Corona Probe
Generating Volt Meter
CPO

## Slide 9
Tandetron electrostatic accelerator
Ion Source Neg potential deck
++++
Stripper
Ion Source Neg potential deck
Accelerate Negative Ions
q Change
Accelerate Positive Ions
SF6 Pressure Vessel
Low Energy Section
High Energy Section
High Voltage Terminal
Gas Stripper
(Recirculating gas
Equal Potential Rings
Ion Source
Injection Magnet/Bouncer
Tank (Pressure Vessel)
Analysing Magnet
Switch Magnet
Motor
Generator
Diode Stack
HVG Driver (PWM 42kHz)
Generating Volt Meter
CPO

## Slide 10: High Voltage Generators
Tandetron Solid-state Charging System
Diagram: http://www.pelletron.com/charging.htm
Pelletron Charging System
Generating Volt Meter (GVM)
Pelletron Chain

## Slide 11: Monitoring and regulating the terminal potential
Generating Volt Meter (GVM), Old ANTARES Version
How do we measure absolute millions of volts?
Generating Volt Meter (GVM)

## Slide 12: GVM Operation
C
Capacitance
Dielectric Constant
Area of Plates
Distance between plates
C
Capacitance
Charge
Voltage
i
Current
Charge
time
Depth Representing
Voltage
Terminal
Tank Wall
Sulphur Hexafluoride (6 Bar)
Capacitance

## Slide 13: GVM Operation
M
Unipolar DC Signal (Approx. 20Hz Bandwidth), is calibrated and displayed as the absolute terminal voltage, and provided on many systems as a feedback to the power generator or Terminal Potential Stabiliser (TPS).

## Slide 14: CPO Operation
How do we measure rate of change?
As with other Terminal sensing devices, the Capacitive Pick Off CPO plate is lacked in the pressure vessel wall of the accelerator, and sits opposite the terminal. It consists of a simple disc. Though many are supplied as dual inputs which were intended to cancel mechanical motion, all our accelerators only use 1 channel.
The CPO compliments the slower GVM signal sensing the terminal ripple, and provides this signal as an input of the Terminal Potential Stabiliser (TPS)
Depth Representing
Voltage
Capacitance

## Slide 15: Corona Probe
How to remove chare from the terminal?
The Corona Probe is located on the wall of the pressure vessel, featuring pointed needle which are driven to an appropriate location to allow the emission of elections from the ground plane, through the gas to the terminal as a regulated corona discharge by a voltage controlled constant current source, emitting electrons from the tips of the needles. Typical NEC systems have 0-200 uA units operating typically between 5-20 uA.
N. R. Lobanov, M. C. Blacksell, P. Linardakis, and D. Tsifakis, “Terminal Voltage Stabilization of Pelletron Tandem Accelerator”, in Proc. HIAT'12, Chicago, IL, USA, Jun. 2012, paper TUB03, pp. 124-128.

## Slide 16: Corona Probe
The Corona Probe can be adjusted from its initial position at the pressure vessel wall with adjustable linear movement reducing or increasing the gap between the needles and the high voltage terminal.
The emission of electron from the needles on the probe are contingent on their sharpness, and the strength of the electric field to emit electrons from the needle points.
The gap or distance between the needles and the terminal has to be adjust to be in the “goldilocks” zone as to have an electric field high enough to emit electrons, but not too high as to course external breakdowns, or damage to the corona probe cable.
Terminal
Terminal
Volts/mm
Consideration

## Slide 17: Corona Probe (V to I Regulator)

## Slide 18: Corona Probe (V to I Regulator)
VACUUM TUBE, HIGH VOLTAGE BEAM TRIODE RCA/GE/SYLV #6BK4C (OR EQUIV)
To prevent an electron flow from the cathode to the plate the control grid is driven negative to suppress the electrons back to wards the cathode.
To allow a current to flow the grid is allowed to go more positive to allow electrons the transvers the tube to the plate.
The Control Grid Voltage is a key parameter to observe when adjusting and accelerator running conditions.
Note: due to the high voltage between the plate and the cathode, electrons can gain enough energy to generate some x-rays.

## Slide 19: Liner
A liner is an insulated plate sitting opposite to the terminal which can push and pull the terminal capacitively driven by a bi-polar HV amplifier
The liner is suggested to be able to correct the terminal bandwidth of 100Hz

## Slide 20: Regulation of Terminal Voltage
Stabilisation or Regulation of the terminal voltage is critical to the desire path the ions go through beam optic components such a magnets and electrostatic analysers which may also ensure that delivered beams to end-stations or detectors is of the correct energy and mass.
Terminal voltage fluctuation can lead to the beam moving left and right within magnetic bends, and not arrive on target, or within the physical window constraints of detectors etc.
There are typical 2 methods or levers in which the terminal voltage can be corrected;
Corona Loading + [Liner]
Change up charge to Terminal

## Slide 21: Terminal Potential Stabiliser
Terminal Potential Stabiliser (TPS)
Corona Prob Style Machines, Sirius and ANTARES

## Slide 22: Conservation of charge
Corona Current
Set Level
Ion Beam Current
Column Current
Leakage Current
Charging Current
Lost Current
Terminal Potential

## Slide 23: Conservation of charge
Corona Current
Set Level
Ion Beam Current
Column Current
Leakage Current
Charging Current
Lost Current
Terminal Potential

## Slide 24: Conservation of charge
Corona Current
Set Level
Ion Beam Current
Column Current
Leakage Current
Charging Current
Lost Current
Terminal Potential
Example:
Chart: 6MV Sirius Total Current Dissipation
| Corona | Lost Charge | Beam | Column
 | 12 | 2 | 1 | 85

## Slide 25: Voltage/Charging Controller
Charge Controller
Charge Control systems used on Star and Vega
Note: Vega Example

## Slide 26: So what else can we do?
Firstly are our feedback signal responsive enough, and can we improve on these?
  Can the GVM provide a faster signal?
Corona control appears as been demonstrated to have inherent response delay in current transfer time.
Can another active element be used such as the Stripper Tube be modulated?
  What would a high speed telemetry system look like to achieve this?

## Slide 27: Thanks
Don’t be strangers, feel free to visit us as ANSTO Lucas Heights 

## Slide 28: Acknowledgement Accelerator Systems Development Team
dbu@ansto.gov.au
David Button
Philip Chatfield
Andrew Downes
David Garton
John Lawlor
Yasser Mafi Nejad
Michael Mann
Sam Micallef
Nikolas Paneras
Matthew Rees
Craig Thompson
Jian Wang