## Slide 1: The southern hemisphere’s first X-band radio-frequency test facility at the University of Melbourne.
Matteo Volpi on behalf of the group: Roger Rassool, Suzie L. Sheehy, Geoffrey Taylor, Scott David Williams (The University of Melbourne), Rohan Dowd, (AS - ANSTO)
ATF-IC-CoP 2022

## Slide 2
X-band Laboratory for Accelerators and Beams (X-LAB)
  A new laboratory will be created at the University of Melbourne (UoM) to build a next generation compact light source.
    First high-power, high-frequency accelerator laboratory in the South Hemisphere.
    Testing high gradient structure prototype and RF components for CLIC
    Ultra-precision manufacturing
    Design and develop more widely available high quality x-ray sources
  This project will provide local researchers and students with an opportunity to make significant advances in accelerator design
  Excellent potential for applications in medical radiation therapy and beyond

## Slide 3: Location of Xband Facilities and High Gradient Tests Stands
Half of a CERN high-gradient test facility has been sent to Melbourne to serve as the central infrastructure for further development in the new X-Lab.

## Slide 4: Long Journey
CERN
Melbourne
CERN news !
On 16 September 2020, a container filled with pallets, boxes and electronic racks left CERN’s Meyrin site to embark on a two-month sea journey to the other side of the world.

## Slide 5: Melbourne Bunker a few months ago
Previously used for 35 MeV betatron
~70m2: Space for x-band linac and medical applications
  Refurbishing scheduled for the next months

## Slide 6: Melbourne Bunker a few days ago
Previously used for 35 MeV betatron
~70m2: Space for x-band linac and medical applications
  Refurbishing scheduled for the next months

## Slide 7
At the heart of the project is the X-band high power equipment including modulator and klystron (left) feeding rf power to two copper accelerating structures (right). An electron source will also enable acceleration of electron beams.

## Slide 8: Mel-Box Bunker test stands
CERN
UoM
Schematic of the X-lab setup for experimental structure testing and conditioning CLIC prototype structures. During phase 2 the electron source will be added to create an electron accelerator.
Phase I
Control room

## Slide 9: MeL-BOX Layout
x6 FPGA IQ demodulation 400-MHz AC (228MSPS ADCs)Phase and amplitude needed for reliable line/phase switching.
x6:3 FPGA IQ demodulation 400-MHz AC (1600MSPS ADCs)  Most interesting signals for BD physics.
x4:2 Faraday Cups (250MSPS ADCs)  Used to interlock the system and are interesting for high gradient physics.
x6 Log detectors ~50MHz Bandwidth, 45dB dynamic range (250MSPS ADCs)  Used for interlocking the system.
DUT
DUT
FPGA IQ demodulation slow ADCs
FPGA IQ demodulation fast ADCs
Faraday cups
Log detectors

## Slide 10: MeL-BOX main components
NI-PXI Crate
TD 24 Structure
Pulse Compressors
Vacuum Rack
3D titanium load prints
Interlock Crate
LLRF Rack

## Slide 11: Conditioning CLIC prototype structures (TD24)
CLIC structures high power tested so far.
Facilities need to be able to predict breakdown to ensure reliability of operation while in collider applications the large number of rf structures also require reduced likelihood of breakdown. Each RF breakdown causes a destructive sideways kick to the particle bunches. This is why accelerating structures are conditioned - or ‘broken in’ - before operation. The process of conditioning requires a variation of input pulses as gradient and pulse width to ensure reliability of operation
to attain the lowest possible breakdown rate.
(ns)
TD24 structure so far reached 40.45 MW (97.9 MV/m)

## Slide 12
X-band for Industrial and medical technologies
Phase II
See Scott W. talk: Simulations and design of a compact beamline at the University of Melbourne X-lab

## Slide 13
Local Manufacturing
Australian National Fabrication Facility (ANFF): manufactured a W90 to waveguide adaptor
  “This project has been a good chance to review some of our processes and equipment gaps.”
They have been using monocrystal diamond tooling
  surface roughness ~Ra 7nm (best measured so far)
  tolerances within 1µm
Be able to make disc? Not sure, but that our overall aim for the future is to have that capability.

## Slide 14: Structure production flow

## Slide 15: Dark current and radiation in high-gradient RF structures
High surface electric fields lead to spontaneous emission of electrons from the material surfaces.
Field emitted electrons can be captured by the EM fields and be accelerated throughout the structures, causing the so-called dark currents.
If these electrons become energetic enough they can also produce ionizing radiation when colliding with material walls.
Curtesy of David Caballero
e- energy
TD24 structure

## Slide 16: Radiation Simulation
A simulation of the bunker area has been performed with FLUKA software. The simulation includes:
  Two high gradient structures with upstream, downstream and later electron sources
  Bunker geometry including old electron beam holes and emergency exit (weak point)
    The bunker area was hosting a 35 MeV electron beam betatron
  Lead shielding around the structures:
    Downstream lead bricks around the faraday cup
    Two lead walls
  20 cm thick chicane wall plus interlock gate.
20MeV e-
Downstream
Faraday Cup
1MeV e-
Upstream
0.5MeV e-
Structure
Bunker
Structures
Bunker Geometry
Emergency exit
z
x
Beam line
x
y
z
x

## Slide 17: DOSE-EQ
Soil
Soil
Dark Matter Lab
DOSE-EQ above the annual threshold 1mSv
DOSE-EQ > 1mSv/year
No limits applied
Setup Summary:
  20 MeV
  100ns flat top
  0.5 mA dark current
  200 Hz repetition rate
  2000 working hours per year, (40h/week x 50 weeks)
  Shields:
    Lead bricks around downstream faraday cup
    Wall bricks around the structures

## Slide 18: Conclusion and future plan
A multi-million-dollar package of X-band equipment has been sent to Melbourne to serve as the central infrastructure for further development in the new X-Lab.
This project will shift Australia’s long-standing position as an excellent user of particle accelerators, to a world-leading developer of this cross-cutting technology.
Refurbishing of the Uni. basement is done at 90%:
  Renew the radiation protection license
  X-LAB philanthropic funding by Baker foundation
  LIEF grant for beamline diagnostics
  First high power Xband pulses, beginning 2022
  Welcome to visit the facility!
    Help and support are more than welcome !