## Slide 1: Operation status and future plans of the SRF module at NSRRC
National Synchrotron Radiation Research Center
RF Group
Zong-Kai Liu
2023/04/12
Asian Forum for Accelerators and Detectors 2023 Melbourne, Australia, April 12-14 2023

## Slide 2
Two Light Sources in NSRRC
TLS Storage Ring (1.5 GeV)
25 beam lines
TPS Storage Ring (3.0 GeV)
Phase I + Phase II: 16 beam lines
Phase III: 9 beam lines (~2026)
Office & Lab
Administration Building

## Slide 3
Main parameters | TLS | TPS
Circumference [m] | 120 | 518.4
Beam energy E [GeV] | 1.5 | 3.0
LINAC [MeV] | 50 | 150
Beam current [mA] | 360 | 500
Natural emittance εx [nm-rad] | 22 | 1.6
RF frequency [MHz] | 499.654 | 499.654
Harmonic number | 200 | 864
Momentum compaction | 6.78 x 10-3 | 2.4 x 10-4
Energy spread σE/E | 7.56 x 10-4 | 8.86 x 10-4
Energy loss/turn (dipole) [keV] | 128 | 852.7
RF voltage [MV] | 1.3 | 3.0 (1.5x2)
RF Plant | CESR-type SRF module | KEKB-type SRF module
Straight sections [m] | 6m X6 | 12m X6 & 7m X18
Machine Parameters

## Slide 4
TLS RF System (Storage Ring)
CESR type SRF Module
100 kW klystron transmitter
Analog LLRF

## Slide 5
Operation Status of TLS SRF System

## Slide 6
Operation Status of TLS SRF System
Recent major SRF trip events－unknown quench：
  Vacuum activity in RBT/FBT，but there is no vacuum event in the window.
  LHe pressure does not change.
  Pt falls rapidly、Pt/Tuner phase changes rapidly、Pr rises rapidly within ~ 12 us.
Partial warm up (to 60 K) and run with lower RF voltage (from 1.6 MV to 1.3 MV).
Vacuum
Pr
Pt/Pf phase
Pt
It may be the arc/MP/other activity in the cavity, which may be related to the heavy gas loading.

## Slide 7
Operation Status of TLS SRF System
Recent major SRF trip events－Vacuum burst with tuning angle change：
  During the LHe refilling of a SC wiggler (SW) at the down stream of SRF module.
  There is always a 22 Hz beam tube vibration in Z direction, and resulting in a oscillation of tuning angle.
  Occurred frequently during 2011~2013 with 1.6 MV operation (solved by optimizing the LHe filling process of SW and applying beam processing/aging).
  Repeated several times in 2022~2023 with 1.3 MV operation, no vacuum activity during aging/processing. (the original filling process is changed to reduce heat load of SW)
Window CCG
Tuning angle
Beam Tube Vibration

## Slide 8
TPS RF System
300 kW transmitter & LLRF
KEKB type SRF module
300 kW klystron
SRF electronics
Two RF plant in operation
KEKB type SRF module
300 kW klystron transmitter for 1 station
Analog LLRF

## Slide 9
Operation Status of TPS SRF System

## Slide 10
Operation Status of TPS SRF System
Trip during early stage of operation：
  Due to Multipacting (MP) in the coupler
  It may continuously cause trip
  Apply bias voltage (+1000 ~ +1200) to avoid MP

## Slide 11
LINE Message
API
LINE
Notify
RF soft alarm
Accelerator operation alarm
SRF trip indicator
RF Trip analysis
Oscilloscope display
Real time signal inquiry (check message machine status)
Display on cell phone/PC
LINE-based RF Soft alarm System
Soft alarm system is developed to shorten the downtime and diagnosis time、as well as to improve the reliability (avoid trip due to slow changed signals).
All messages and figures are sent to user’s smart phone and PC via LINE (an instant messaging app.)

## Slide 12
Future Plan (1)－Power Combine
Vc#3 1.42 MV
Ib 500 mA
Pr#3 10 kW
Pf#3 245 kW
Vc#2 1.42 MV
Pr#2 12 kW
Pf#2 243 kW
Stored current reached 500 mA at the end of 2020.
Several insertion devices will be installed into the storage ring.
RF power of 2 stations with 300 kW klystron may reach the limit in 3 years.
500 mA | Total Pb | 2 x RF
w/o ID | 427 kW | 213 kW
Phase I | 476 kW | 238 kW
Phase I+II | 526 kW | 263 kW
Phase I+II+III | 576 kW | 288 kW
Beam power only and with all ID gap close to the minimum.
One solution is: Combination of an 80-kW SSPA to a 300-kW klystron in TPS.

## Slide 13
Future Plan (1)－Power Combine
Combine method was tested in the RF Lab.

## Slide 14
Future Plan (2)－SC Harmonic Cavity
Bunch lengthening by the passive 3rd harmonic cavities: to reduce the beam-induced heat on magnet arrays of IDs and to improve the beam lifetime.
1.5 GHz, Two-cell SRF cavity designed by NSRRC. (2016-2018)
Niobium cavity contracted to MHI MS, Japan. (2020-2021)
EPs and annealing to be done at KEK. (2023)
Cryostat to be designed and assembled by NSRRC. (2022-2023)
HOM absorbers contracted to RI, Germany. (2022-2024)
Assembly and integration at NSRRC. (2024-2025)

## Slide 15
Future Plan (2)－SC Harmonic Cavity
Courtesy of Ming-Chyuan Lin (chyuan@nsrrc.org.tw)
Courtesy of Chih-Hung Lo (lo.ch@nsrrc.org.tw)
HOM absorbers of panel type:

## Slide 16: Summary
Two light sources (TLS & TPS) use 500 MHz SRF modules as their accelerating cavity, MTBF of both RF systems are all > 500 hours in 2022.
Major trip types and improvements in recent years:
  CESR type for 1.5 GeV TLS
    Unknown quench events: Partial warm up and lower operating RF voltage.
  KEKB type for 3.0 GeV TPS
    Multipacting in coaxial coupler: applying positive bias voltage.
    Develop soft alarm system to shorten the downtime and improve the reliability.
Future plans for TPS SRF activity:
  Power combination for each SRF module.
  SC 3rd harmonic cavity to lengthen the electron bunch.

## Slide 17
Thank you