## Slide 1
(The project is jointly supported by IUAC and DAE-BRNS)
Dr. Subhendu Ghosh
(on behalf of all the extended group members of FEL and all the national/international collaborators)
Inter University Accelerator Centre, New Delhi, India
Commissioning status of Compact FEL-THz facility at IUAC and the first energy measurement of the electron beam

## Slide 2
Plan of Presentation
Outline of the Compact FEL facility at IUAC (Delhi Light Source)
First production of dark current and photocurrent electron beam
Ongoing commissioning status of the various sub-systems
Conclusion

## Slide 3
Plan of Presentation
Outline of the Compact FEL facility at IUAC (Delhi Light Source)
First production of dark current and photocurrent electron beam
Ongoing commissioning status of the various sub-systems
Conclusion

## Slide 4
Bunch length ~ 3 mm, E-energy = 7 MeV, t  10 ps onwards
=wavelength of radiation
7
1
6
5
3
2
S
Now all the electron disks emit radiation in synchronism, & the
light can amplify itself to form high-intensity laser radiation.
t0-t
t0
t0+t
7
1
6
5
3
2
S
1
2
3
5
6
7
Interaction of Photon and wiggling electron inside undulator magnet
U – Undulator wavelength
BU – Undulator mag field
Outline of the Compact FEL facility at IUAC
Compact FEL facility at IUAC – A Pre-bunched Free Electron Laser : How it is different from the conventional FEL?
Microbunching
in FEL
(Osc., Seeded and SASE)
CONVENTIONAL FREE ELECTRON LASER

## Slide 5
Starting point of Delhi Light Source
Conventional FEL
Prebunched FEL (DLS)
How can we produce
Micro-bunched
(‘Comb’) beam ?
Outline of the Compact FEL facility at IUAC
Compact FEL facility at IUAC – A Pre-bunched Free Electron Laser : How it is different from the conventional FEL?

## Slide 6
10 Hz
0.375 fs – 6.6 ps
An electron gun – laser operated PC & a resonator powered by klystron/modulator
A laser system – Ultra-short laser pulses – produces ultrashort electron bunches
– single ultrashort laser pulse is split into many (2,4,8,16)
3. Photocathode – Copper and Cs2Te (Deposition system)
Solenoid, other electromagnets – focus/steer electron beam – Cavity to Undulator
5. Undulator – to produce THz radiation
6. Beam diagnostic and e.m. radiation detector systems
7. Electronics and Control system
U – Undulator wavelength
BU – Undulator mag field
Outline of the Compact FEL facility at IUAC
Major components
Electron
beam
Laser
beam
Solenoid
Solenoid
Laser
Undulator Magnet

## Slide 7
5 MHz
0.333-6.6 ps
U – Undulator wavelength
BU – Undulator mag field
f ~ 3 – 0.18 THz
T = 333 fs to 6.6 ps
 = 100 m to 2 mm
Laser pulses as well
Electron bunches
½ w plate rotates S-wave by 45o.
Polarizing beam splitter makes S-wave & P-wave by reflection & transmission.
16 micro bunched laser within 6-7 psec.
Tuning of THz frequency
Outline of the Compact FEL facility at IUAC

## Slide 8
~ 3 μs (RF width)
160ms (6.25 Hz)
200 ns
Total no. of laser
pulses (5 MHz)
are 15
333 fs (3 THz)
200 ns
~ 6 ps
So total no. of laser micropulses/e-bunches/THz pulses 15  16  6.25 = 1500 pulses/sec
Scheme of Production of electron beam microbunches - Multi-bunch train of THz
3 s
3 s – nominal RF pulse width
6.25 Hz – Klystron/Mod freq.
Laser Osc. Frequency – 5 MHz
Laser splitting freq. – 3 / 0.18 THz
1
2
15
1
15
1
16
Outline of the Compact FEL facility at IUAC

## Slide 9
1
16
6 ps
1
2
349.65 ps
8580
1
16
6 ps
3 s
572
1
16
6 ps
200 ns
Resonator freq. = 2860 MHz, (T = 349.65ps), Laser freq. = 5 MHz, (T = 200 ns)
RF pulse duration = 3 sec, Rep rate of RF = 3.125, 6.25 Hz up to 50 Hz
Acceleration of multi-micro bunch train by proper phase of RF
Outline of the Compact FEL facility at IUAC

## Slide 10
FEL Beam Line
FEL Hall –
Class 10000
Clean room
THz line
Electron line
THz Expt. chambers
3
2
1
Outline of the Compact FEL facility at IUAC

## Slide 11
Outline of the Compact FEL facility at IUAC

## Slide 12
Frequency | Wavelength | Energy of the THz pulse | Time width of the THz pulse train | Power of THz in one pulse train
(THz) | (mm) | (μJ) | (ps) | (kW)
0.18 | 1.66 | 0.044 | 189.3 | 0.232
0.6 | 0.5 | 0.7 | 60 | 11.6
1 | 0.3 | 2.88 | 40 | 72
1.5 | 0.2 | 1.88 | 26.6 | 71
2 | 0.15 | 1.75 | 20 | 87.5
3 | 0.1 | 1.15 | 16 | 72
Expected Deliverables from FEL facility of IUAC
Time width | No. of electrons | Total electron current
~ 300 fs | 9.3  107 | 75 A
~ 6 ps | 1.5  109 | 40 A
~ 3 s | 2.25  1010 | 1.2 mA
1 sec. | 1.4  1012 | 22.5 nA
Electron beam
THz Radiation
Outline of the Compact FEL facility at IUAC

## Slide 13
Plan of Presentation
Outline of the Compact FEL facility at IUAC (Delhi Light Source)
First production of dark current and photocurrent electron beam
Ongoing commissioning of the various sub-systems
Conclusion

## Slide 14
Electron
beam
Laser
beam
Solenoid
Solenoid
Laser
YAG
+
Beam
Catcher
First Photocathode electron beam on YAG screen
and beam Catcher (08 Feb 2022)
First production of dark current and photocurrent electron beam
The e-gun was only conditioned up to 1 MW
in absence of High Power Circulator

## Slide 15
Bending Magnet #1
BPM#4
BPM#3
e- beam direction
Input Power of RF gun (MW) | Acc. Field of RF gun (MV/m) | Mag. Field (BM#1) (Gauss) | Measured Energy of e-beam (MeV)
3 | 56.5 | 500 | 4.0
First energy (4.0 MeV) measurement of the electron beam from RF Photocathode gun of IUAC
Results (13 Feb. ’23)
13 Feb. 2023
BPM#3 (before BM)
BPM#4 (after BM)
Photoelectron
High Power Circulator
was installed in July 2022

## Slide 16
Bending Magnet #1
e- beam direction
Input Power of RF gun (MW) | Acc. Field of RF gun (MV/m) | Mag. Field (BM#1) (Gauss) | Measured Energy of e-beam (MeV)
4 | 65 | 550 | 4.5
More energy (4.5 MeV) measurement of the electron beam
10 March 2023
BPM#1
BPM#2
BPM#3
BPM#4
12 | ~109 |  | 8
8 MeV is expected within a few months
Results (10 March ’23)

## Slide 17
Achievements:
Photo-electron is produced and transported up to the Expt. Station after B.M.#1
Electron Energy is measured, already 56% of the designed value is achieved
A pilot experiment with the electron beam is being planned
Undulator is successfully characterized at IUAC and is fully operational
Demonstration of the first signature of THz is expected soon (by closing the jaws)
First production of dark current and photocurrent electron beam

## Slide 18
300
1
16
6 ps
Present limitations:
Presently nano-second laser pulse is producing the electron and occupies 3 RF period
No Phase synchronization between nano-second laser pulse with RF, Electron beam has long energy tail
The femto-second laser device will solve this
Installation of the femto-second laser system is expected during Oct-Dec. 2023
First production of dark current and photocurrent electron beam

## Slide 19
Plan of Presentation
Outline of the Compact FEL facility at IUAC (Delhi Light Source)
First production of dark current and photocurrent electron beam
Ongoing commissioning status of the various sub-systems
Conclusion

## Slide 20
High power RF System of DLS
Chronological developments:
RF conditioning was accomplished w load at 25 MW, 4µs, 50 Hz (2019)
Vacuum Circulator could not be installed. Proposed SF6 circulator.
SF6 based Circulator could be installed only in July 2022.
Conditioning of e-gun could be started beyond 1 MW after July ‘22
Present power reached 5 MW (4 µs), 8.5 MW (1.2 µs)
Consists of solid state modulator based Klystron
Peak Power = 25 MW, 4µs pulse, 50Hz rep. rate
Vacuum based wave guide section
Specifications:
Klystron & Modulator
Waveguide connection
Inside the FEL hall
Waveguide & e-gun
connection
Status in 2019
Status in
July 2022
SF6 based
Circulator
Ongoing commissioning status of the various sub-systems
A DUMMY WAVEGUIDE SECTION WAS CONNECTED
IN PLACE OF VACUUM CIRCULATOR

## Slide 21
Ongoing commissioning status of the various sub-systems
Photocathode deposition
system
Electron Gun
BPM#1
BPM#2
Undulator
BPM#3
BPM#4
BM#1
ICT#1
ICT#2
Steerer

## Slide 22
Dipole, Quadrupole and Steering magnets
2 BMs, 7 Quads, 5 steering magnets
are designed and developed at
Electromagnetic Applications
and Instrumentation Division,
BARC in record time.
Thanks to Shri Sanjay Malhotra & his team, BARC
P.S. of all the magnets - IUAC
Magnets arrived at IUAC
on 24 Jan. 2020
Ongoing commissioning status of the various sub-systems

## Slide 23
Initially metal (Cu) photocathode is being used for electron beam generation.
In later stage, semiconductor photocathode will be used for higher beam current.
Polished surface
Cu-Be Spring for RF contact
Present status:
One Cu PC was inserted into the RF cavity and RF conditioning has been started.
First electron beam from Cu photocathode has been demonstrated.
Deposition system for Cs2Te photocathode is commissioned, deposition will start soon
Cathode plug during inserting into RF gun
Ongoing commissioning status of the various sub-systems
DLS: Photocathode (Cu – Initial, Cs2Te - final)

## Slide 24
Photocathode Deposition System designed by IUAC, fabricated at BNL, USA, with participation of IUAC.
The system was arrived at IUAC in July’21. It was commissioned and first deposition is expected – by 2023.
Salient features of PC deposition System:
System consists of 4 XHV chambers inter-connected to each other.
Operating vacuum ~ 10-11 mbar.
Capable to develop mono and/or multi-alkali advanced photocathodes
by co-evaporation process.
Capable of (a) in-situ thickness monitoring,
(b) Q.E. measurement during the deposition process
(c) heating and cooling arrangements of the substrate etc.
Movement/transfer of the cathode plug from one chamber to another.
Temperature calibration on Mo plug is presently going on
CsTe deposition is expected to be started in next few months
Ongoing commissioning status of the various sub-systems

## Slide 25
PC Insertion Chamber
Storage Chamber
Magnetic manipulators
Deposition Chamber
Initial cleaning Chamber
Ongoing commissioning status of the various sub-systems
Photocathode Deposition System

## Slide 26
Arrival and Commissioning of the Undulator (donated by HZB (BESSY), Refurbished at DESY) at IUAC
Undulator (U48) along with its other accessories reached IUAC on 25th December, 2020.
Current Status: Field measurement was done at IUAC (reported in InPAC2022),
Was commissioned and aligned in the beam line. It is ready to produce THz
Undulator
Ongoing commissioning status of the various sub-systems

## Slide 27
Final positioning of the magnet is done in 2023 after alignment/installation of vacuum chamber
Vacuum chamber alignment
Linear Stages
Alignment Fixture
Horizontal Alignment bolt
Outer dimensions:
15.2 mm (Hor.) x 40 mm (vert.)
Inner dimensions:
14 mm (Hor.) x 25 mm (Vert.)
Length ~ 2500 mm
Alignment of U48
Using Distance between Taylor-Hobson Fiducials and Magnet Centre (Beamline Height)
Wall reference
U48 in beam line
Ongoing commissioning status of the various sub-systems

## Slide 28
266nm laser pulse with time width ~ 200-300fs, Initial goal ~ 500 fs
Laser pulse energy at 266nm ~ 10 J, Initial goal is 1.5 J
Splitting the laser pulse into 2,4 and farther maximum upto 16 micropulses, Initial goal is 4 micropulses
Stable synchronization with the master clock which drive the electron gun
Electron gun is powered with 2860 MHz master clock with a 4 microsecond RF window. The laser pulses should be present inside this RF window with proper synchronization of master clock.
The Fiber laser system of DLS is currently being developed in collaboration with KEK, Japan with the active participation of three young personnel of IUAC under the supervision of RRCAT personnel
Case | Copper PC | Cs2Te PC
Charge/bunch (pC) | 3 | 15
No. electrons/bunch | 9.38e+07 | 9.38e+07
Q.E. of PC | 1.0 e-05 | 1.0 e-03
No. photons required | 9.38 e+12 | 9.38 e+10
Photon energy required | 1.5 J | 72.6 nJ
The major parameters
Ongoing commissioning status of the various sub-systems

## Slide 29
Feedback
(electronics) for Synchronization between laser and RFM or PCB
Laser Master (2860MHz or 1300MHz)
For IUAC DLS
Block diagram:
Fibre laser
)Oscillator(Yb-doped fiber)
1030nm (wavelength)
~130 mW (ave. power)
~25 nm (bandwidth)
Pulse duration: 1.2 ps
Pre amplifier
(Yb-doped PCF)
1030nm (wavelength)
CW pumping
Bandwidth 15 nm
pulse duration 30 ps (expected)
Pulse energy 25 nJ
Stretcher
(Gratings)
50% energy loss
bandwidth)15 nm
Burst Amp I
(Yb doped PCF rod)
1030nm (wavelength)
Gain 260X
3.13 Hz rep rate
Laser Micro-Bencher
1030nm (wavelength)
3.13 (rep.rate)
1-15pulses@5MHz,
1 - 16 pulses @ THz
Compressor(Gratings)
1030nm (wavelength)
Energy loss maximum 30%
4HG
(Telescope-BBO-BBO)
258 nm (wavelength)
3.13 Hz (rep.rate), 1-15pulses@5MHz,
1 - 16 pulses @ THz
Individual pulse energy is achieved so far 0.1 uJ
LTL+FF
(UV telescope)
258 nm (wavelength)
3.13 Hz (rep.rate)
~250 fs (pulse duration)
RF Sync.
With laser clock
Sync. Trigger (12.5Hz)
Sync. RF REF. (5 MHz and 130 MHz )
Phase locked loop
Burst Amp II
(Yb doped PCF rod)
1030nm (wavelength)
Gain 260X
3.13 Hz rep rate
Individual pulse energy 8.5 uJ
SOA pulse picker
5-10 MHz from 130 MHz in 4 µs window
Pre-pre amplifier Yb doped fiber
Gain 30X
1030 nm wavelength
Bandwidth 25 nm
Pulse duration 2ps
B
A
D
E
F
G
H
I
C
J
L
K
M
N
Ongoing commissioning status of the various sub-systems

## Slide 30
Oscillator
Pulse picker
Pre amplifier
Burst amplifier I and II
Compressor and FHG
Pre-pre amplifier and stretcher
Schematic of the
fiber laser system
Ongoing commissioning status of the various sub-systems

## Slide 31
Oscillator 130 MHz
Pre Amplifier
SOA pulse picker 130 MHz to 5/10 MHz
Present status
Oscillator.+Pulse Picker+Pre. Amplifier
Burst amplifier stage 1 and 2
Lay out
Pulse picker output
FHG 266nm

## Slide 32
Oscillator 130 MHz
Pre Amplifier
SOA pulse picker 130 MHz to 5/10 MHz

## Slide 33
Fiber Laser System
Present status:
FHG is demonstrated
Pulse width ~ few ps
Pulse Energy ~ tens of nJ
Hardware for splitting of a single laser in to 2,4, … is ready
Future Tasks:
Pulse width is to be reduced to a few hundred of femto seconds
Output energy per pulse to be enhanced ~ 1.5 J
Splitting of a single pulse in to 2,4,…. is to be demonstrated (proven technique)
Upgradation of a few existing components is needed.
Ongoing commissioning status of the various sub-systems

## Slide 34
Plan of Presentation
Outline of the Compact FEL facility at IUAC (Delhi Light Source)
First production of dark current and photocurrent electron beam
Ongoing commissioning status of the various sub-systems
Conclusion

## Slide 35
Conclusion
Delhi Light Source - at the stage of final commissioning
Photocurrent electron beam is produced and 4.5 MeV is measured, Routine production ~ an hour
Need to condition e-gun up to 15 MW, then increase the electron energy from 4.5 to 8.0 MeV
Essential BL components and the undulator are installed and functional
Pilot experiment with electron beam will be performed soon
Commission/operate femto-second fiber laser at IUAC to produce monoenergetic, ultrashort e-bunches
Signature of THz by undulator should be demonstrated: by end of 2023 (after installation of fibre laser)

## Slide 36
Acknowledgement
BRNS-DAE – partial funding for the project
Core members of FEL group at IUAC
Associated members of FEL group at IUAC
BARC (Shri Sanjay Malhotra),
RRCAT (Dr. P.K.Mukhopadhyay, Dr. K.K.Pant),
VECC (Dr. Vaishali Naik)
Dr. Triveni Rao, BNL, USA
Prof. Junji Urukawa, Dr. Alexander Aryshev, KEK, Japan
Dr. Markus Tischer, DESY, Germany
Dr. Peter Mitchel, HZDR, Germany
Dr. D.Kanjilal, Dr. Amit Roy, Former Director, IUAC
Dr. R.K.Bhandari, Honorary Visiting Scientist, IUAC
Prof. A.C.Pandey, Director, IUAC

## Slide 37
Thank you for your kind attention

## Slide 39
𝑇𝑝𝑢𝑙𝑠𝑒 = 𝜆𝑟(𝑁𝑢 + 𝑁𝑚 ∕𝛽) ∕𝑐
Deliverables: THz radiation
Frequency | Wavelength | K | Qtotal, Nm | Bavg | E(real) | T-pulse | P(real) | P_avg @ 5MHz laser rep rate and @ 6 Hz RF of 3 μS
(THz) | (mm) |  | (pC, no.) |  | (μJ) | (ps) | (kW) | (μW)
0.18 | 1.666666667 | 3 | 30,2 | 0.45 | 0.044 | 189.3 | 0.232 | 3.96
0.6 | 0.5 | 2.43 | 60,4 | 0.77 | 0.7 | 60 | 11.6 | 63
1 | 0.3 | 1.2 | 120,8 | 0.72 | 2.88 | 40 | 72 | 259
1.5 | 0.2 | 1 | 120,8 | 0.52 | 1.88 | 26.6 | 71 | 169
2 | 0.15 | 0.99 | 120,8 | 0.43 | 1.75 | 20 | 87.5 | 157
3 | 0.1 | 0.6 | 240,16 | 0.2 | 1.15 | 16 | 72 | 103.5

## Slide 40
Electron
beam
Laser
beam
Solenoid
Solenoid
Laser
E-gun – accelerated field was created during Feb-March 2021 (<1 MW)
First e-beam (dark current) was observed on 01 March 2021
Electron was pulled out from the surface, accelerated and
focused by solenoid and reached the YAG screen.
YAG
Electron beam (Dark Current)
on YAG Screen
Sometime, the dark current can be very intense
First production of dark current and photocurrent electron beam

## Slide 41
Steerer
Quadrupole
Installation of BARC made electromagnets in the beam line of DLS
BM#1
Steerer