## Slide 1
Perspectives on compact and
next-generation particle therapy systems
Suzie Sheehy
Associate Professor and Baker/ANSTO Fellow
University of Melbourne
suzie.sheehy@unimelb.edu.au
@suziesheehy
12/4/2023

## Slide 2: UniMelb Particle Therapy Research Topics
  1. Smaller? Next-generation synchrotrons
    Smaller: superconducting synchrotrons and magnets
    NIMMS: “Next ion medical machine study” w. CERN
    PhDs: Frank Zhang, Hannah Norman
  2. Faster? Beam Delivery Systems
  The lack of ability to change energy rapidly is a major bottleneck in duration and quality of treatment. Focus on Beam Delivery System.
  - Dr. Jacinta Yap w. Adam Steinberg, Hannah Norman (PhDs)
  3. Better? Novel Fixed-Field Accelerators:
    Superconducting FFA hadron therapy designs (Adam Steinberg) to meet future paradigms (FLASH, multi-ion etc)
MedAustron, Austria

## Slide 3: Next generation facilities & advanced CPT delivery techniques…
FLASH: Shoot-through, single field, uncertainties with motion, ELST long timescale
Arc therapy: Complex delivery, long BDT
Multi-ions: Range of momenta for different ions, prohibitively high B fields needed
MR-guided PBT: Field corrections, physical constraints with BDS
Voltz et al., Front. Oncol., Volume 12 - 2022 | https://doi.org/10.3389/fonc.2022.930850

## Slide 4
Smaller.
Greg Pieris
Me.
Frank Zhang

## Slide 5: Medical Superconducting Synchrotron
Goal: Design a smaller heavy ion therapy synchrotron for next gen. facilities
NIMMS Superconducting Compact Synchrotron, courtesy of E. Benedetto
What makes this design special?
5
Based on 90° CCT magnets at 3.5 T
27 m circumference, 1.89 m bending radius
Combined-function quadrupole layers inside dipole
Adapt magnets from synchrotron to gantry
Perform beam dynamics simulations:
Study unique magnet effects
Slide thanks to: Hannah Norman

## Slide 6: Medical Technology
Gantries (LBNL, HUST) (in progress)
LBNL (top), HUST (bottom) SC gantry layouts
Parameter | LBNL | HUST
B[T] | 2.4 | 2.46
Gradient [T/m] | 17.43 | 12.0
Current [A] | 922 | 584.38 (dip), 370.84 (quad)
Physical Bend [°] | 67.5 | 90
Slide thanks to: Hannah Norman
Notes: One of the things to note about these magnets is that they can have multiple layers of magnets, so instead of being separated, we can layer over the bending, focusing and other functions…

## Slide 7: Automated Optimisation Algorithm for Compact Synchrotron Design
X. Zhang, S. Sheehy, DOI: 10.1063/5.0047815
Slide thanks to: Frank Zhang
Frank Zhang

## Slide 8
Faster.
Sienna Clarke (USyd)
Adam
Jacinta

## Slide 9: Treatment delivery
J. Yap et al., “Future Developments in Charged Particle Therapy: Improving Beam Delivery for Efficiency and Efficacy,” Front. Oncol., vol. 11, pp. 1–25, 2021.
(A) Accelerator
(B) Beam delivery system (BDS)
(C) Beam delivery methodology
+
+
=
Treatment delivery
Slide thanks to:
Dr. Jacinta Yap
Notes: Look at how a treatment is delivered in terms of the tech
- simple overview: source, main transport & delivery technology and the method of delivery
- In proton therapy, modern methods: scanning many pencil beams at maximum depth of the tumour in one layer, then decrease energy and scan again repeatedly for full coverage of the tumour volume

## Slide 10: How can we improve the delivery of CPT?
Total treatment time =
  Patient related activities (immobilisation, imaging, unload etc.)
  +
  Equipment related checks (couch positioning, beam checks etc.)
  +
  BDS: Beam delivery time (BDT)
    Beam-on: Irradiation
    Beam-off: Requested, waiting
Beam delivery process:
  Beam production (A)
  Beam transport (B)
  Dose delivery (C)
  Corrective/Mitigation approaches
Patient setup
Beam delivery
Equipment checks
Total treatment time
Faster treatment means better:
1. Availability: Cost, throughput
Delivery efficiency
2. Utilisation: Effectiveness, quality
  Treatment efficacy
Slide thanks to:
Dr. Jacinta Yap
Notes: Many things can be done… any improvement in treatment efficiency is valuable
Beam delivery process is defined by the BDS & its capabilities... if we can improve this then doubly beneficial

## Slide 11: Beam delivery time: a bottleneck
The constraint which determines the BDT is the energy layer switching time (ELST)
ELST is the most dominant contributor to BDT at 71%
BDS efficiency directly related to utility & availability, 70-90% of BDT for majority of tumours <1 L
  1 min reduction in BDT could result in 10x (prostate) patients/day
J. Yap et al., “Future Developments in Charged Particle Therapy: Improving Beam Delivery for Efficiency and Efficacy,” Front. Oncol., vol. 11, pp. 1–25, 2021.
(K. Suzuki et al., 2016, Med. Phys., vol. 43, doi:10.1118/1.4952731)
(J. Shen et al., 2017, Med. Phys., vol. 44, doi:10.1002/mp.12504)
(K. Suzuki et al., 2011, Med. Phys., vol. 38, doi:10.1118/1.3604153)
Slide thanks to:
Dr. Jacinta Yap
BDS: Biggest driver of delivery efficiency & treatment efficacy
Notes: First looking at efficiency
Looking at the scanning delivery process, the longest wait is actually for your machine and BDS to shift between energies
This contributes to the largest time component when your beam is on and running
Different machines have different speeds... Increasing time penalty for longer ELSTs on BDT, clear bottleneck

## Slide 12: ELST reduction on treatment quality
Treatment planning optimisation: number of spots & energy slices for required dose coverage
Does minimising the number of layers affect the dosimetric quality of plans?
Numerous studies show significant reductions to BDT & little effect on treatment conformity
  Kang et al. PMB (2008), Cao et al. PMB (2014), Li et al. Int. J. Radiat. Oncol. Biol. Phys. (2015), Muller et al. PMB (2016)
TPS optimisation: reduction of ELSTs decreasing BDT by 40% without compromise in treatment quality
  “﻿Delivery times of intensity modulated proton therapy can be reduced substantially without compromising robust plan quality. Shorter delivery times are likely to reduce treatment uncertainties and costs.”
Decreasing ELSTs can improve dose uniformity and homogeneity (lung treatments)
  “﻿The interplay effect is highly patient specific, depending on the motion amplitude, tumor location and the delivery parameters of the field.”
(S. Dowdell et al., 2013, PMB., vol. 58, doi:10.1088/0031-9155/58/12/4137)
(S. Van De Water et al., 2015, Int. J. Radiat. Oncol. Biol. Phys., vol. 92, no. 2, pp. 460–468.)
Slide thanks to:
Dr. Jacinta Yap
Notes: So a faster delivery can lead to faster treatment, more patients but what about treatment quality/efficacy?
Actually shown that even if you use diff optimisation approaches and use different ways to reduce ELST, you can substantially reduce BDT with no cost to treatment quality

## Slide 13: Why so slow?
Shorter ELST & BDT: benefits both throughput & treatment quality
Multiple avenues to improve the BDT -> directly with the BDS?
Typical beamlines/gantries: momentum acceptance ±1%, energy acceptance ±2%
  Magnet settings need to be readjusted for a change of ~3-5mm depth
  Change magnetic fields to transport particles with different energies
What if we could deliver multiple energies without changing the beamline magnets?
Large energy acceptance (LEA) system accommodates beams of varying energies with same optics
Novel implementation of fixed field alternating gradient (FFA) optics
No waiting time on magnets: eliminates/minimises the ELST
TURBO: Technology for Ultra Rapid Beam Operation
Yap et al.. “Preliminary study of a large energy acceptance FFA beam delivery system for particle therapy, IPAC2021
Slide thanks to:
Dr. Jacinta Yap
Notes: How can we do this? One way is by building an improved BDS.
So we know that a faster ELST leads to a shorter BDT
Minimise this means we could effectively deliver multiple energies and therefore layers rapidly
Scan in multidirectional forward and back, in depth as rapid as transverse
adjustment of typically 5mm requires the entire beamline to be reset/retuned to accommodate for diff beam energies
Ensure reproducibility of beam characteristics. Beam position, characteristics have to be within all margins at all rotations

## Slide 14
Scaled-down BDS prototype is being developed at the University of Melbourne
  Low energy (0.5-3.5 MeV) protons & He, DC Pelletron accelerator
First demonstration of hadrons transported through FFA optics
Explore feasibility as a fixed-beamline solution for CPT
Modularity & flexibility to iterate for rapid lab implementation
(Courtesy of A. Steinberg)
Slide thanks to:
Dr. Jacinta Yap
Project TURBO: Technology for Ultra Rapid Beam Operation
Notes: Demonstrator being built at UniMelb
Adapted for our low energy acc, to explore and de-risk these concepts
Unique environment lab

## Slide 15: TURBO Research Path
1. Treatment Planning Studies
2. Scaled-down Prototype
3. Full-Scale Design
2. (Today’s focus)
What beam characteristics are necessary?
  Beam characterisation & simulation studies
How do we vary the energy rapidly (and confirm it)?
  Rapid variable energy beam: Degrader studies
How do we transport all energies simultaneously?
  FFA arc lattice design and magnets
Integration studies:
  Matching magnets, diagnostics, control system etc
Rapid scanning system and dosimetry

## Slide 16: TURBO integration studies on Pelletron
3D Printed magnets
6
Slide thanks to:
Adam Steinberg

## Slide 17: The Pelletron
Electrostatic accelerator (DC)
Protons up to 3.5MeV
  Other ions may be available
Beam currents 0.1 – 3.5µA
Beam sizes ~1mm (collimated)
Multiple beamlines (3 currently)
  Rotates between beamlines
8.85m
4
Slide thanks to:
Adam Steinberg
Notes: Pelletron lab serves many projects, but crucially, they are all coarse (they want the beam for what it can do, but don’t care that much about the beam properties themselves). Their experiments are for ion implantation or detector checking, where a reasonably uniform beam can be shot at a target for a long time – but that’s not our goal.

## Slide 18: The Pelletron: Beam Characterisation
5
What does the beam look like?
Phase Space
Position vs angle at some point
Energy Distribution
Add energy-dependent dispersion
Slide thanks to:
Adam Steinberg
Notes: As well as making clear that our beamline will attach to the Pelletron somewhere (in one of those spaces), our beamline design is very different to those already there (quite obviously: ours must be bendy, to disperse out the orbits at different energies, whereas all the others are straight in order to get maximum beam to various experiments).
Our needs are different to those of the work with the Pelletron, so we need to do experiments to investigate the machine to check that our plans actually will work.

## Slide 19: The Pelletron Lab: Our Equipment
200 µm
For Phase Space
UniBEaM Fibre Scanner
Tantalum Sheets
7
Slide thanks to:
Adam Steinberg
Notes: As well as making clear that our beamline will attach to the Pelletron somewhere (in one of those spaces), we should state explicitly that our beamline design is very different to those already there (quite obviously: ours must be bendy, to disperse out the orbits at different energies, whereas all the others are straight in order to get maximum beam to various experiments).
Our needs are different to those of the work with the Pelletron, so we need to do experiments to investigate the machine to check that our plans actually will work.

## Slide 20: Results To-Date: Phase Space
Example Measurement:
With projections on position/angle axes
Energy
Vary
Particle
Time
Results Overlayed:
Centred and with approximate ellipse shown
12
Slide thanks to:
Adam Steinberg

## Slide 21: The Pelletron Lab: Our Equipment
For Energy Degrader
UniBEaM Fibre Scanner
25.4 µm thick
Kapton Sheets
4 3 2 1
Dipole Magnet
8
Slide thanks to:
Adam Steinberg
Notes: As well as making clear that our beamline will attach to the Pelletron somewhere (in one of those spaces), we should state explicitly that our beamline design is very different to those already there (quite obviously: ours must be bendy, to disperse out the orbits at different energies, whereas all the others are straight in order to get maximum beam to various experiments).
Our needs are different to those of the work with the Pelletron, so we need to do experiments to investigate the machine to check that our plans actually will work.

## Slide 22: Results To-Date: Energy Degrader
3MeV Proton Beam:
Before and After dipole kick
After one Kapton Sheet:
Simulation agrees with Experiment
Distribution is too broad to measure spectrum, but sufficient to confirm mean beam energy
Kick smaller than 3MeV case due to limited beampipe size
11
Slide thanks to:
Adam Steinberg
Sienna Clarke (USyd)

## Slide 23: FFA Arc Design (where the magic is…)
STEP 1: design an arc where orbits change with energy
Triplet cells (FDF)
Rectangular magnets
Up to decapole order
No need to super fine-tune
STEP 2: Optimiser Outputs
Dispersion much reduced everywhere
Residual dispersion way smaller
Tune variation much reduced
Beam stable for all energies
Slide thanks to:
Adam Steinberg
Note: in accelerator terms, this reaches the same ‘range’ of energies as a clinical system would need… ie. a “scaled” system.

## Slide 24
11
Slide thanks to:
Adam Steinberg

## Slide 25: Beam tracking: how are the beam spots?
Slide thanks to:
Adam Steinberg
Current work:
Error studies (looks OK)
Magnet designs to Increase ‘realism’ of simulations

## Slide 26: BDS Scaling up for the clinic
Fast delivery (reduction of treatment time)
  Cost/throughput, rescanning, moving organs
  Be able to treat for more indications (i.e. energy variation with motion)
Deliver beam with position (isocentre) & spot size (field size) within required margins
  Required beam characteristics & unaffected by varying the magnets
Economical (size, weight, costs)
  Footprint, operational, shielding
Yap et al.. “Preliminary study of a large energy acceptance FFA beam delivery system for particle therapy, IPAC2021
Questions for a Fixed-beamline solution
Momentum range: energy bands, particle type?
Design requirements: magnets, nozzle, beam parameters?
Clinical acceptability & utility: workflow, specialised treatment sites?
Collaboration with CERN (NIMMS) on electromagnetic model to study and adapt for TURBO
Slide thanks to: Hannah Norman
Notes: Several questions and considerations to think about with, want to ensure what we do remains within the realm of possibility for treatment: TURBO technology translated to the clinic
Lots of opportunities to explore, particularly for fixed line delivery and with the progress in this field, it's clear that emerging techniques will require better BDS technology with better capabilities

## Slide 27: TURBO Research Questions:
TURBO BDS demonstrator for scaled-down conditions being developed, driven by clinical requirements:
1. Treatment Planning Studies
2. Scaled-down Prototype
3. Full-Scale Design
2.
What beam characteristics are necessary?
  Beam characterisation & simulation studies
How do we vary the energy rapidly (and confirm it)?
  Rapid variable energy beam: Degrader studies
How do we transport all energies simultaneously?
  FFA arc lattice design and magnets
Integration studies:
  Matching magnets, diagnostics, control system
Rapid scanning system and dosimetry:
  This part (mostly) is commercially available

## Slide 28
Thanks!
Dr. Suzie Sheehy
Suzie.Sheehy@unimelb.edu.au
@suziesheehy

## Slide 29: Scaling Up for TURBO
Full-scale clinical beamline requires:
  Higher strength magnets (maximum field ~3.5 T):
    treat in modes or energy regions
  Good field quality
  Compactness
Curved Canted-cosine-theta (CCT) arrangement may be suitable:
  Strong focusing
  Large field to accommodate heavier ions to full kinetic energies
  Compact and lightweight
Collaboration with CERN (NIMMS) on EM model to study and adapt for TURBO
Alternating-Gradient (AG) Quadrupole layers 1,2
Dipole layers 3,4
Alternating-Gradient CCT dipole magnet, adapted from W. Wan et al, DOI: 10.1103/PhysRevSTAB.18.103501
Slide thanks to: Hannah Norman

## Slide 30: Fields experienced by particles along beamline

## Slide 31: TURBO Research Questions:
TURBO BDS demonstrator for scaled-down conditions being developed, driven by clinical requirements:
1. Treatment Planning Studies
2. Scaled-down Prototype
3. Full-Scale Design
2.
What beam characteristics are necessary?
  Beam characterisation & simulation studies
How do we vary the energy rapidly (and confirm it)?
  Rapid variable energy beam: Degrader studies
How do we transport all energies simultaneously?
  FFA arc lattice design and magnets
Integration studies:
  Matching magnets, diagnostics, control system
Rapid scanning system and dosimetry:
  This part (mostly) is commercially available

## Slide 32: Final thoughts… MR-guided PT?
“A magnetic decoupling of the PBS system and the MRI scanner […] remains indispensable”
In this experimental study, effects of the beamline (rather than nozzle) magnets were investigated… and it was concluded that:
“No visual image degradation was observed due to the static settings of the beam line magnets.” [emphasis mine]
Sebastian Gantz et al 2020 Phys. Med. Biol. 65 215014
So in future, we might STILL be limited in energy steps, and only change layers every 2s?!

## Slide 33: Some potential questions re. MR-PT
What level of background magnetic field from the transport beamline/gantry is acceptable, or can be shielded (dynamically or passively)?
What dynamic variations of external magnetic field (dB/dt) from the transport beamline/gantry are acceptable?
Could MR-PT be better – if the energy layer switching time were faster? (i.e. 4D dynamic beam delivery)
NB. in some literature ‘4D delivery’ means varying the beam intensity in time.

## Slide 34: Australian Strategy & Roadmap
The Australian Collaboration for Accelerator Science (ACAS)
Australia is entering a new era for accelerator development and applications, including new activity in hadron therapy, space technologies, and advanced materials development.
These developments are supported by the skills and experience of the existing community, but for longevity and high impact outcomes of these areas there needs to be a long-term roadmap for accelerators and the pipeline of related skills and expertise. ACAS is a joint initiative between ANSTO’s Centre for Accelerator Science and ANSTO Australian Synchrotron, ANU’s Heavy Ion Accelerator Facility, and the University of Melbourne.
ACAS will bring together the wider accelerator community across Australia - both accelerator technologists and accelerator users - to collaborate and to pool skills, resources, and networks as a critical mass for securing opportunities and influence on the national and international level.
It will provide a platform for coordination and collaboration of Australia’s activities and developments in accelerator science and its applications, including next generation machines for future light sources, ion beams and microprobes, medical and industrial accelerators, as well as high energy particle and nuclear physics.
- Excerpt from 1-page submission to NCRIS Roadmap

## Slide 35: The Pelletron Lab: Our Setup
1445 mm
500 mm
9
Slide thanks to:
Adam Steinberg

## Slide 36: Results To-Date: First Light
What went wrong and
how did we solve it?
High noise in all cases
  Removed other Pelletron diagnostics
  Increased measurement time
  Smoothing if required
High signal liable to saturation
  Reduced beam current
  Attenuated to reduce intensity
10
Slide thanks to:
Adam Steinberg

## Slide 37: What is a disruptive technology?
Clayton Christensen introduced the idea of disruptive technologies in a 1995 Harvard Business Review article*.
“sustaining technologies tend to maintain a rate of improvement; that is, they give customers something more or better in the attributes they already value”
“Disruptive technologies introduce a different attributes from the mainstream … and often perform far worse along one or two dimensions that are particularly important to those customers.
As a rule, mainstream customers are unwilling to use a disruptive product in applications they know and understand.
At first, disruptive technologies tend to be used and valued only in new markets or new applications; in fact, they generally make possible the emergence of new markets.
*https://hbr.org/1995/01/disruptive-technologies-catching-the-wave

## Slide 38: Disruptive technologies
Innovation/technology that:
“displaces an established technology and shakes up the industry”
or
“a ground-breaking product that creates a completely new industry”
Notes: Often, the performance of the disruptive technology is often LESS than that of existing mainstream technology.