## Slide 1: Ways to collect CT data sets on IMBL
Anton Maksimenko

## Slide 2: Basic concepts
detector
motors
shutter
IOC
IOC
IOC
Commander station:
Experiment control, logic, GUIs, user’s scripts
EPICS channel access
Data storage
MASSIVE HPC
User’s equipment
Optical link
User’s PC
Rendering station:
Rendered views, processing
SSH, virtual desktop
samba

## Slide 3: CTgui
When you start CTgui it restores previous settings.

## Slide 4: Configuration
When you start the data acquisition all parameters are automatically saved in the working directory in the acquisition[.number].configuration file.
[General]
description=Energy: 81keV\nObjective to detector 525mm\nMax pixel size 7.65 micron per pixel
workingdir=/home/imbl/temp
syncdetectordir=false
acquisitionmode=On-the-fly (software)
doserialscans=true
doflatfield=true
dodyno=false
domulti=false
prerun=
postrun=
[serial]
endCondition=number
number=3
time=@Variant(\0\0\0\xf\0\0\x3\xe8)
script=
ongoingseries=false
ffoneachscan=false
scandelay=@Variant(\0\0\0\xf\0\0\0\0)
serialmotor=SR08ID01SST24:SAMPLEV
serialstep=1
serialirregularstep=false
[scan]
motor=SR08ID01SST24:ROTATION
click

## Slide 5: Logging
The data collection is logged into acquisition.[number].log file.
2015-11-04_11:51:08.283653 Acquisition started with the image filename prefix "DF_Z0_BEFORE".
2015-11-04_11:51:08.947105 0 0.001
2015-11-04_11:51:09.072739 1 0.001
…
2015-11-04_11:51:11.434004 19 0.001
2015-11-04_11:51:11.434074 Acquisition finished.
2015-11-04_11:51:16.456792 Acquisition started with the image filename prefix "BG_Z0_BEFORE".
2015-11-04_11:51:17.081726 0 0.001
2015-11-04_11:51:17.213108 1 0.001
…
2015-11-04_11:51:19.669598 19 0.001
2015-11-04_11:51:19.669624 Acquisition finished.
2015-11-04_11:51:23.873773 Acquisition started with the image filename prefix "SAMPLE_Z0_T".
2015-11-04_11:51:24.550976 0 0.298
2015-11-04_11:51:24.777280 1 0.588
2015-11-04_11:51:24.812643 2 0.588
2015-11-04_11:51:24.983257 3 0.687
2015-11-04_11:51:25.186077 4 0.924
2015-11-04_11:51:25.246472 5 0.924
2015-11-04_11:51:25.397815 6 1
2015-11-04_11:51:25.599135 7 1.209
… 2015-11-04_11:55:51.797736 1850 185.45
2015-11-04_11:55:51.797773 Acquisition finished.
2015-11-04_11:58:50.348342 Acquisition started with the image filename prefix "SAMPLE_Z1_T".
2015-11-04_11:58:50.998557 0 0.408
2015-11-04_11:58:51.153026 1 0.501
… 2015-11-04_12:06:35.443674 Acquisition finished.
Time stamp | Counter | Rotation position | 
2015-11-04_11:51:08.283653 | Acquisition started with the image filename prefix "DF_Z0_BEFORE". |  | 
2015-11-04_11:51:08.947105 | 0 | 0.001 | 
2015-11-04_11:51:09.072739 | 1 | 0.001 | 
… | … | … | 
2015-11-04_11:51:11.434004 | 19 | 0.001 | 
2015-11-04_11:51:11.434074 | Acquisition finished. |  | 
2015-11-04_11:51:16.456792 | Acquisition started with the image filename prefix "BG_Z0_BEFORE". |  | 
2015-11-04_11:51:17.081726 | 0 | 0.001 | 
2015-11-04_11:51:17.213108 | 1 | 0.001 | 
… | … | … | 
2015-11-04_11:51:19.669598 | 19 | 0.001 | 
2015-11-04_11:51:19.669624 | Acquisition finished. |  | 
2015-11-04_11:51:23.873773 | Acquisition started with the image filename prefix "SAMPLE_Z0_T". |  | 
2015-11-04_11:51:24.550976 | 0 | 0.298 | 
2015-11-04_11:51:24.777280 | 1 | 0.588 | 
2015-11-04_11:51:24.812643 | 2 | 0.588 | 
2015-11-04_11:51:24.983257 | 3 | 0.687 | 
2015-11-04_11:51:25.186077 | 4 | 0.924 | 
2015-11-04_11:51:25.246472 | 5 | 0.924 | 
2015-11-04_11:51:25.397815 | 6 | 1 | 
2015-11-04_11:51:25.599135 | 7 | 1.209 | 
… | … | … | 
2015-11-04_11:55:51.797736 | 1850 | 185.45 | 
2015-11-04_11:55:51.797773 | Acquisition finished. |  | 
2015-11-04_11:58:50.348342 | Acquisition started with the image filename prefix "SAMPLE_Z1_T". |  | 
2015-11-04_11:58:50.998557 | 0 | 0.408 | 
2015-11-04_11:58:51.153026 | 1 | 0.501 | 
… | … | … | 
2015-11-04_12:06:35.443674 | Acquisition finished. |  | 

## Slide 6: CTgui – data saving
Any pre-existing data in the directory are overwritten.
Warning!

## Slide 7: Error reporting and debugging
Red labels indicate an impossible parameter and prevent execution.
There is something wrong in two tabs
There are problems
with two parameters
click
Start is disabled
Limit is active

## Slide 8: Acquisition modes
Stop-and-shot
On the fly – software implementation
On the fly – hardware triggered

## Slide 9: Flat field acquisition
On-the-fly mode
Stop-and-shot mode

## Slide 10: Serial scans
Scan counter
  Regular step or manual positioning
Time counter
Conditional completion

## Slide 11: Imaging techniques
Available only in step-and-shot mode

## Slide 12: Dynamic shot

## Slide 13: Tiles

## Slide 14: File naming
Stop-and-shot template
<type>[_Z<series>]_T<projection>[_L<loop>][_S<sub-loop>][_<image>].tif
Examples: SAMPLE_T0278_L02_S1.tif, BG_Z02_T035_5.tif
On-the fly template
<type>[_Z<series>][_T<projection>][_<before/after>[_<image>]].tif
Examples: SAMPLE_Z1_T567.tif, DF_AFTER_05.tif

## Slide 15: Expandability
Script execution is available before and after
Whole acquisition procedure
Each independent scan
Each loop in the tiled mode
Each frame

## Slide 16: Thank you