## Slide 1: Neutron imaging instrument DINGO – moving to 1μm resolution by U. Garbe

## Slide 2: DINGO team and supporter
Team:
Floriana Salvemin
(Joseph Bevitt)
Supporter and consultant:
Eberhard Lehmann
Burghard Schillinger
Mohamed Arif
ISNR-members
Technical Support:
David Roach
Andrew Berry
Frank Darmann

## Slide 3: New Imaging Beamline at ANSTO DINGO Layout
Primary Collimator
Selector Wheel / Secondary Shutter Unit
Flight Tubes
Tertiary and Fast Shutter
Rail System
Sample Stage
Detector Stage
Beam Stop
Sliding door

## Slide 4
Major components
Shielding elements build

## Slide 5
Selector wheel / shutter unit
Internal structure filled with steel / wax mixture
Modular system
Different inserts to generate different beam sizes at sample position
One block design
Adjustment screws for fine alignment

## Slide 6
Selector wheel
Shutter unit

## Slide 7
Instrument characteristics
Sample stage
Specs translation:
400mm travel length (xyz)
~ 500kg load
Specs rotation stage 1:
360° rotation stage
resolution of 0.001°
~200kg load
Specs rotation stage 1:
resolution of 0.001°
~20kg load
Detector box
Ikon-l CCD, NEO CMOS camera
Two Zeiss macro lens (50mm and 100mm)
Three beam sizes 200 x 200, 100 x 100 and 50 x 50 mm2
Pixel size 10 – 100 µm
25fps fast imaging under development

## Slide 8
Resolution (L/D) | Dingo - peak foil [n/s*cm2] | Dingo - foil average [n/s*cm2]
500 | 5.59 * 107 | 5.30 * 107
1000 | 1.06 * 107 | 9.00 * 106
ratio | 5.3 | 5.9
DINGO flux distribution

## Slide 9: Neutron tomography measurement
[E. Lehmann, PSI]

## Slide 10: Attenuation Coefficient (X-rays)
[A. Kaestner, PSI]

## Slide 11: Attenuation Coefficient (neutrons)
[A. Kaestner, PSI]

## Slide 12: History of Concrete
First concrete experience approx. 2 million years old
1400BC in Greece
Ancient roman infrastructure

## Slide 13: Concrete world record
The world record for the largest concrete pour in a single project is the Three Gorges Dam in Hubei Province, China by the Three Gorges Corporation. The amount of concrete used in the construction of the dam is estimated at 16 million cubic meters over 17 years. The previous record was 12.3 million cubic meters held by Itaipu hydropower station in Brazil.

## Slide 14: Economic importance of concrete
Globally consuming 25 gigatonnes/year
annual costs of remediation and premature replacement reaching 3.8% of global GDP
Reinforced Concrete is the most widely used building material, owing to its versatility and relatively low cost
If we make Australian roads last 5 years longer it would pay off for hundreds of OPAL research reactors

## Slide 15: Porosity distribution in concrete sample

## Slide 16
Visualising enamel, apatite,
border in transparent mode
Visualising enamel transparent
and apatite solid
Visualising root canal in
transparent grey
In collaboration with
International Centre for Theoretical Physics in Trieste (Italy)
Neutron imaging on ancient teeth

## Slide 17
Treated and Untreated Tooth

## Slide 18
Ceramic on left side removed

## Slide 19
Enamel on left side removed

## Slide 20
Dentin on left side removed

## Slide 21
Metal and filler material of upper tooth only

## Slide 22
Filler material of upper tooth only

## Slide 23
Dingo radiography – “Woods metal”
Wt%
Bi – 42.5
Pb – 37.7
Sn – 11.3
Cd – 8.5
Non-eutectic
Solidification Processing – 25-28 July 2017, SP17

## Slide 24
Dingo radiography – “Woods metal”
0.1C/min
Solidification Processing – 25-28 July 2017, SP17

## Slide 25
First results of 14μm pixel size tomography
Transparent setting to visualize internal defects
Cut through reconstructed sample
Setup:
10μm Gadox screen with IKON-L CCD
ZEISS macro lens with Kenko adapter
Peng, E et al, RSC Adv. 7(43), 27128-27138 (2017)

## Slide 26
Composite Manufacturing Robot

## Slide 27: Porosity analysis after CT reconstruction

## Slide 28: Comparison of total volume of porosity

## Slide 29: Void development after deformation

## Slide 30: Comparison of porosity level between the undeformed and deformed samples
Chart: HGT temperature (°C)
| 950.0 | 850.0 | 750.0 | 650.0
Deformed Samples | 0.043 | 0.03654 | 0.088434 | 0.75017
 | 0.0045 | 0.0196 | 0.031 | 0.3238

## Slide 31: DINGO upgrade with PSI neutron microscope

## Slide 32: New scintillation screen to deliver 2.5μm pixel size
Isotope enriched 157Gd2O2S:Tb
Porous Si-membrane
Thinner screen increase light output
First results with 2.5μm pixel size lead to 5μm optical resolution

## Slide 33: Neutron Microscope
Magnetic imaging
Neutron microscope
Standard neutron imaging
In-situ imaging under extreme conditions
Resolution
Light element imaging
Aircraft parts
Ammonites
Welds
Cartriges
Zr-hydride
Bio-membranes
Nanotubes
Bragg edge scanning
1m
1µm
1nm

## Slide 34: Instrument layout
70-150m
3-4m
Magnification of ~ 20-50

## Slide 35: First results
Magnification factor 1
Magnification factor 5
Simulated image with low and homogeneous neutron beam divergence
First experiment using VCN SANS setup:
Cd plate with 3 apertures
(being f = 1 mm, f = 2 mm circles and a 0.5 mm wide slit)
Simulated image with high and inhomogeneous neutron beam divergence (VCN)

## Slide 36: Neutron microscope summary
Sextupole lens works as magnifier
Magnification factor 5 proven by experiment
New instrument with 70 – 100 m length on cold source could achieve magnification of 50
100-50nm resolution depends on detector
Potential resolution of 0.2 - 0.5nm limited by the wavelength of cold neutrons

## Slide 37
Summary
Thank you!
Please contact us : ulf.garbe@ansto.gov.au

## Slide 38
World Conference on Neutron Radiography (WCNR-11)
02-07 September 2018
Maritime Museum Sydney, Australia
Notes: ANSTO and Sydney looks forward to the Board’s decision so we can make the 2018 meeting a truly memorable and highly successful event which will create a long-lasting legacy for ISNR.