## Slide 1: Development of event-type neutron imaging detectors at the energy-resolved neutron imaging system RADEN at J-PARC
Joe Parker
CROSS/BL22 Group
WCNR-11 Australian National Maritime Museum, Sydney, Australia 4 September 2018
Notes: Good morning. My name is Joe Parker.
Thank you very much for having me here today and thank you for allowing me speak to you about my research on ‘A new micro-pattern detector for high-rate, time-resolved neutron imaging and its applications’.

## Slide 2: RADEN and µNID development members
JAEA/J-PARC Center Takenao Shinohara
Tetsuya Kai
Kenichi Oikawa (BL10)
Masahide Harada (BL10)
Takeshi Nakatani
Mariko Segawa
Kosuke Hiroi (BL15)
Yuhua Su (SE Group)
CROSS Hirotoshi Hayashida
Joe Parker (µNID Lead Developer)
Yoshihiro Matsumoto
Nagoya University Yoshiaki Kiyanagi
Kyoto University Toru Tanimori
Atsushi Takada
WCNR-11 4 Sep 2018 J. Parker

## Slide 3: J-PARC (Japan Proton Accelerator Research Complex)
LINAC
(330m)
3GeV Synchrotron
(25Hz,1MW)
Neutrino facility
(to Kamiokande)
Hadron experimental facility
Materials and Life science experimental Facility (MLF)
Nuclear transmutation facility
50GeV main ring
(0.75MW)
Tokyo
J-PARC
Materials and Life science experimental Facility (MLF)
100 km from Narita international airport
MLF is currently operating at 500kW
(successful test at 1 MW in June 2018)
WCNR-11 4 Sep 2018 J. Parker
Notes: 1 day by walk, 2 or 3 hours by a bus or train

## Slide 4: RADEN/BL22 – Neutron imaging instrument at the MLF
World-class, pioneering instrument for pulsed-neutron imaging
Leading facility for conventional radiography in Japan
Commissioning from Nov. 2014, user program from April 2015
World’s first instrument dedicated to energy-resolved neutron imaging using pulsed neutrons
First images from RADEN (7 Nov 2014)
WCNR-11 4 Sep 2018 J. Parker

## Slide 5: Neutron imaging detectors at RADEN
Camera type
Event type
µNID
nGEM
LiTA12
Mirror
Scintillator
CCD
camera
Neutron color I.I.
Single-mirror CCD system (Andor iKon-L, EMCCD)
Neutron color I.I. (high-res,high-speed)
Micropattern detectors (nGEM, µNID)
Li-glass pixel scintillator (LiTA12)
micropattern
micropattern
Li-glass scintillator
WCNR-11 4 Sep 2018 J. Parker

## Slide 6: ERNI with event-type detectors
Energy-dependence → quantitative information on macroscopic distribution of microscopic quantities
Pulsed neutrons → wide energy range, accurate energy determination by time-of-flight
Requires detector with:
  Sub-mm spatial resolution
  Sub-µs time resolution
  Mcps count rate
  Strong background rejection
Energy-dependent neutron transmission
Energy
1 meV
1 keV
Wavelength
10 Å
10-2 Å
Resonance absorption
Bragg-edge,
Magnetic imaging
WCNR-11 4 Sep 2018 J. Parker

## Slide 7: Event-type detectors at RADEN
nGEM
GEM surface
170 µm
µNID
Anode
structure
5 cm
LiTA12
10 cm
10 cm
 | nGEM | µNID | LiTA12
Detector type | Micropattern | Micropattern | Scintillator
Converter material | 10B | 3He | 6Li
Active area | 10 cm x 10 cm | 10 cm x 10 cm | 5 cm x 5 cm
Spatial resolution | 1 mm | 0.1 mm | 3 mm
Time resolution | 10 ns | 250 ns | 40 ns
Efficiency (thermal) | 10% | 26% | 23%
Count rate | 0.2 Mcps | 1 Mcps | 6 Mcps
Gamma sensitivity | 10-4 | < 10-12 | low
WCNR-11 4 Sep 2018 J. Parker
Notes: 30 sec

## Slide 8: Current detector performance
G O A L
Chart: Spatial resolution (mm)
| 0.2 | 1.0 | 3.0
Y の値 1 | 0.4 | 0.3 | 6
nGEM
LiTA12
Count rate and spatial resolution should be considered together
µNID
WCNR-11 4 Sep 2018 J. Parker

## Slide 9: Development of event-type detectors
For energy-resolved neutron imaging, we need:
  Improved rate performance (10 Mcps order or higher)
  Improved spatial resolution (µNID <100µm, LiTA12 <1mm)
  Full integration into RADEN experiment control system
  Fast, easy-to-use data processing software
Ongoing development of µNID and LiTA12
  Optimization of existing detector components, development of new components
  Optimization of analysis algorithms, measurement methods
  Development of new control and processing software
WCNR-11 4 Sep 2018 J. Parker
Notes: Benefits
  Drastically reduced measurement times
  Complicated measurement techniques
  Combine with CT, real time measurements

## Slide 10: µNID development at RADEN
WCNR-11 4 Sep 2018 J. Parker
Optimization of filling gas, analysis algorithms
Full integration into RADEN control system
Web-based analysis GUI

## Slide 11: µNID development at RADEN
G O A L
Chart: Spatial resolution (mm)
| 0.2 | 1.0 | 3.0
Y の値 1 | 0.4 | 0.3 | 6
nGEM
LiTA12
µNID
µNID
(ca. 2014)
WCNR-11 4 Sep 2018 J. Parker

## Slide 12: µNID: µPIC-based neutron imaging detector
Gaseous time-projection-chamber
  CF4-iC4H10-3He (45:5:50) at 2 atm
  µPIC micropattern readout
  Compact ASIC+FPGA data encoder front-end
3-dimensional tracking of decay pattern + time-over-threshold
  Accurate position reconstruction
  Strong gamma rejection
33 cm
E
TOT for proton-triton track
Proton
Triton
Neutron
Digital encoder with time-over-threshold (TOT)
Threshold
µPIC
Discriminator
Time-over-threshold
(∝ energy dep.)
Neutron detection via 3He
Track length ~5 mm in gas
µPIC readout
10 cm x 10 cm area,
400 µm pitch x,y strips
WCNR-11 4 Sep 2018 J. Parker
Polyimide substrate
Developed at Kyoto University
Notes: Our µPIC-based neutron imaging detector, or µNID, shown here, consists simply of a µPIC, which is just a printed circuit board, and drift cage contained in an aluminum pressure vessel. As stated previously, we achieve neutron detection by adding He3 to the gas of the detector, and we measure both the detailed 3-dimensional track and energy deposition of the resultant proton-triton pairs.
The energy deposition is estimated using the time-over-threshold technique. As illustrated here, it is simply the time for which an analog signal exceeds the threshold of the discriminator. This measurement is accomplished digitally, without the need for ADCs, allowing us to maintain maximal counting rates. In this way, we can measure approximately the energy deposition at each hit point, as illustrated by this sample event.
By observing this TOT distribution, we are able to cleanly separate the proton and triton tracks, leading to a much better estimate of the neutron interaction position, as compared to say taking the midpoint of the track as one would be forced to do in the absence of the TOT information.
This characteristic shape is also what allows us to reject essentially all backgrounds.

## Slide 13: µNID DAQ/control system layout
FPGA-based data encoders and DAQ controller
Data transfer via Gigabit Ethernet
Full integration into RADEN experiment control system
Data encoders
DAQ PC
SiTCP
Encoders
µPIC
Control box
・DAQ control
・system monitor
・DC power
External timing signals
Vessel pressure
Ethernet
GbE × 4
±2.5V, +3.3V
Ambient temperature
Network
DAQ control
Monitoring
Power
Sensor power
~50 cm
WCNR-11 4 Sep 2018 J. Parker

## Slide 14: Spatial resolution at RADEN
0.2
mm
0.3
mm
0.4 mm
0.5 mm
0.1 mm
Transmission
Distance from top (mm)
Bin size: 40 x 40 µm2
8 cm
Optimized gas mixture and refined neutron position reconstruction algorithm
Image of Gd test pattern
  L/D: 5000, Exposure: 1.5 hours
  10% contrast at 5 lp/mm (100µm line width)
Image of Gd test target
WCNR-11 4 Sep 2018 J. Parker

## Slide 15: Hardware rate performance
Revised data encoder hardware (100BASE-T → Gigabit Ethernet)
  About 7 times increase in rate capacity
Changed from Ar to CF4 based gas mixture
  About 2 times increase due to smaller track size
Total increase in global peak rate capacity from 0.6 to 8 Mcps
Usable rate limited by offline analysis
Chart: 100Mb/s vs Gb Ethernet
| 4.0 | 16.0 | 36.0 | 64.0 | 100.0 | 144.0 | 196.0 | 256.0 | 324.0 | 400.0 | 484.0 | 576.0 | 676.0 | 784.0 | 900.0 | 1024.0 | 1156.0 | 1296.0 | 1444.0 | 1600.0 | 1764.0 | 1936.0 | 2116.0 | 2304.0
GbE | 0.013494 | 0.053056 | 0.11635 | 0.20545 | 0.3276 | 0.46853 | 0.64782 | 0.81783 | 1.0544 | 1.2872 | 1.5795 | 1.7992 | 2.0653 | 2.3764 | 2.7077 | 3.0107 | 3.4203 | 3.5605 | 3.6459 | 4.2243 | 4.4966 | 4.5648 | 4.6836 | 4.6438
100Mbps | 0.0081271 | 0.027539 | 0.061581 | 0.114 | 0.18058 | 0.26389 | 0.35759 | 0.46199 | 0.5761 | 0.71552 | 0.86957 | 0.91556 | 0.86262
Chart: Ar vs CF4 based gas mixture
| 4.0 | 16.0 | 36.0 | 64.0 | 100.0 | 144.0 | 196.0 | 256.0 | 324.0 | 400.0 | 484.0 | 576.0 | 676.0 | 784.0 | 900.0 | 1024.0 | 1156.0 | 1296.0 | 1444.0 | 1600.0 | 1764.0 | 1936.0 | 2116.0 | 2304.0 | 2500.0 | 2916.0 | 3364.0 | 3600.0
Ar-C2H6 | 0.013494 | 0.053056 | 0.11635 | 0.20545 | 0.3276 | 0.46853 | 0.64782 | 0.81783 | 1.0544 | 1.2872 | 1.5795 | 1.7992 | 2.0653 | 2.3764 | 2.7077 | 3.0107 | 3.4203 | 3.5605 | 3.6459 | 4.2243 | 4.4966 | 4.5648 | 4.6836 | 4.6438 | 4.7967 | 4.8509 | 4.8369 | 5.0598
CF4-iC4H10 | 0.073673 | 0.15428 | 0.26985 | 0.41907 | 0.60195 | 0.8201 | 1.0706 | 1.3471 | 1.6633 | 2.0053 | 2.3662 | 2.7645 | 3.2055 | 3.6638 | 4.148 | 4.6464 | 5.1692 | 5.7091 | 6.2663 | 6.8572 | 7.3941 | 7.9892 | 8.5706 | 8.4241 | 8.5745 | 8.8597 | 9.2326 | 9.3613 | 9.0537
WCNR-11 4 Sep 2018 J. Parker

## Slide 16: Offline hit clustering algorithm
Raw hits are clustered into neutron events (~14 hits/event with current gas mixture)
  Data taken with fixed area and increasing flux
  Efficiency determined by comparing numbers of raw hits and reconstructed neutron events
Change from single-linkage algorithm to DBSCAN-based algorithm with explicit event pile-up resolution
Event loss improved from 2% at 400 kcps to 2% at 1 Mcps global peak rate
2.5 Mcps
<0.4 Mcps
5.6 Mcps
2.5 Mcps
<1 Mcps
Efficiency
Efficiency
Neutron TOF (ms)
Event reconstruction efficiency
Single-linkage clustering (original)
DBSCAN-based clustering (new)
5.6 Mcps
Typical neutron TOF
WCNR-11 4 Sep 2018 J. Parker

## Slide 17: µNID control software/analysis GUI
New control software
  In use since March 2018
  Based on DAQ middleware
  Full integration into beam line control system
New browser-based UI for offline analysis
  In use since April 2018
  Now preparing first update (simplified interface, better data visualization, etc.)
WCNR-11 4 Sep 2018 J. Parker
µNID analysis GUI
Software frameworks at the MLF
IROHA2 – Experimental device control system with web-based UI (MLF)
DAQ Middleware – Detector control and data collection (KEK)

## Slide 18: Demonstration measurements
Increased rate and integrated control
  Perform complex measurements more easily
Computed tomography with TOF
  Help quantify effects of multiple scattering, beam hardening, etc.
Dynamic samples
  Fold TOF info with motion/process frequency
  Currently limited to cyclical processes
WCNR-11 4 Sep 2018 J. Parker
4 cm
4 cm
K. Hiroi et al., J. Phys.: Conf. Series 862 (2017) 012008
5 cm
Computed tomography
Fe step wedge
Polarization image
Magnetic imaging of running motor
Model electric motor (provided by Hitachi)
2°/step 91 images

## Slide 19: Other µNID development
Small-pitch MEMS µPIC
µNID with Boron converter
WCNR-11 4 Sep 2018 J. Parker

## Slide 20: Small-pitch MEMS µPIC
55 mm
Improve spatial resolution with reduced strip pitch
Develop small-pitch µPIC
  Manufacture using MEMS on silicon substrate (by DaiNippon Printing Company, Ltd.) → uniform structures down to 10 µm
  Successfully produced 215 µm pitch µPIC (down from 400 µm)
  Small (14 x 14 mm2) and larger area MEMS µPICs (55 x 55 mm2) tested at RADEN
100µm
400µm
Cu
10~15µm
4~11µm
15µm
Current PCB µPIC
(400 µm pitch)
MEMS µPIC
WCNR-11 4 Sep 2018 J. Parker
215 µm
Surface of MEMS µPIC
(digital microscope)
215µm MEMS µPIC

## Slide 21: MEMS µPIC tests at RADEN
Sufficient gain confirmed for small-pitch geometry
Imaging confirmed but spatial resolution not improved as expected
Gain instability under neutron exposure → improved by grounding substrate but not eliminated
WCNR-11 4 Sep 2018 J. Parker
Image taken with 215µm pitch MEMS µPIC at BL22
55 mm
Gain stability: new MEMS µPIC with glass substrate prepared
Spatial resolution: optimize gas for shorter tracks
More testing this winter
Cu
+HV
0V
0V
Gain depends on anode/cathode geometry
Chart: Elapsed time (hours)
| 0.0 | 0.00694444444444442 | 0.0138888888888887 | 0.0208333333333334 | 0.0277777777777778 | 0.0347222222222266 | 0.0416666666666666 | 0.0486111111111066 | 0.0555555555555466 | 0.0624999999999866 | 0.0694444444444265 | 0.0763888888888665 | 0.0833333333333067 | 0.0902777777777467 | 0.0972222222221867 | 0.104166666666627 | 0.111111111111067 | 0.118055555555507 | 0.124999999999947 | 0.131944444444387 | 0.138888888888827 | 0.145833333333267 | 0.152777777777707 | 0.159722222222147 | 0.166666666666587 | 0.173611111111027 | 0.180555555555467 | 0.187499999999907 | 0.194444444444347 | 0.201388888888787 | 0.208333333333227
Grounded | 9.8635 | 9.8929 | 9.8986 | 9.9093 | 9.8632 | 9.8815 | 9.8966 | 9.9065 | 9.8947 | 9.8919 | 9.8957 | 9.8911 | 9.8964 | 9.915 | 9.8979 | 9.9166 | 9.9222 | 9.9236 | 9.9209 | 9.9179 | 9.9339 | 9.9008 | 9.9159 | 9.9266 | 9.9224 | 9.9277 | 9.9278 | 9.9327 | 9.9184 | 9.9176 | 9.9286
Floating | 9.9519 | 10.007 | 10.6291 | 10.613 | 10.6689 | 10.4714 | 10.9295 | 11.0746 | 10.7503 | 11.0766 | 10.9433 | 11.2686 | 11.2775 | 11.1215 | 11.0875 | 11.3666 | 11.2097 | 11.2509 | 10.9838 | 11.0157 | 10.5286 | 10.6657 | 10.5826 | 10.6232 | 10.3925 | 10.5038 | 11.0886 | 11.2437 | 11.3161 | 11.4327 | 11.5717
MEMS µPIC gain vs time

## Slide 22: µNID with boron converter (B-µNID)
~5 mm
Al drift cathode (t=1mm)
µPIC readout
10B coating (t=1.2µm)
Increase count rate capacity by reducing event size
  10B (α,Li) for 3x smaller event size than 3He (p,t)
  Trade-off in spatial resolution
µNID with flat boron converter (for initial testing)
  Thin, 1.2µm 10B layer → low efficiency (3~5%)
Need to consider ways to improve detection efficiency
WCNR-11 4 Sep 2018 J. Parker
10cm
10B
(t = 1.2µm)
Expect 20~25 Mcps count rate
and 0.4 ~ 0.5 mm spatial resolution
Boron converter installed in µNID
Underside of drift plane

## Slide 23: Spatial resolution study at RADEN
1.0
0.9
0.8
0.7
0.6
0.5
0.4
Line-width (mm)
Study of spatial resolution, event size vs. gas pressure (1.2 ~ 1.6 atm)
L/D:1000, Exposure time: 15 mins
Spatial resolution estimated from contrast of line-pairs (MTF)
Maximum count rate at hardware limit: 22 Mcps @1.6 atm
Pressure (atm) | 1.2 | 1.4 | 1.6
Average hits/event | 5.86 | 5.42 | 4.82
MTF @0.6mm | 27% | 36% | 41%
Spatial resolution @10% MTF (mm) | 0.50 | 0.48 | 0.45
Gas pressure: 1.4 atm
Bin size: 200 x 200 µm2
Contrast (MTF) from
fit of sine curves
WCNR-11 4 Sep 2018 J. Parker

## Slide 24: Current and expected detector performance
G O A L
Chart: Spatial resolution (mm)
| 0.2 | 1.0 | 3.0
Y の値 1 | 0.4 | 0.3 | 6
nGEM
Li-glass
µNID
Boron µNID
MEMS µPIC
µNID
(current)
WCNR-11 4 Sep 2018 J. Parker

## Slide 25: Li-glass pixel detector
WCNR-11 4 Sep 2018 J. Parker
Improvement of spatial resolution using centroiding, image stepping

## Slide 26: 6Li Time Analyzer, model 2012 (LiTA12)
5 cm
Li-glass
Developed at KEK, Japan
Li-glass scintillator (GS20) pixels on Hamamatsu H9500 multi-anode PMT
3 mm anode pitch limits spatial resolution
Improve spatial resolution using super resolution
  Single, flat scintillator with centroiding
  Image stepping with sub-pixel shifts
WCNR-11 4 Sep 2018 J. Parker
Li-glass detector parameters | 
Area | 5cm x 5cm
Spatial res. | 3 mm
Time res. | 40 ns ~
Eff. (thermal) | 22%
Count rate | 6 Mcps
Image taken at RADEN of cadmium ‘t’
Amp
VME DAQ modules

## Slide 27: Centroiding with single scintillator
Gd indicator (<0.1eV)
Replace pixels with 1mm thick plate
Spatial resolution improved with centroid computation
  ~0.7 mm for Au and Gd
Au indicator (4.9eV)
WCNR-11 4 Sep 2018 J. Parker
T. Kai et al., Physica B, 2017

## Slide 28: Image stepping (multi-image compositing)
LiTA12 (pixel detector) on remote-controlled stage
x,y scan of Gd test chart on 6 x 6 grid
  0.5 mm step size, 36 images
Low resolution image
Simple reconstruction
WCNR-11 4 Sep 2018 J. Parker

## Slide 29: Current and expected detector performance
G O A L
Chart: Spatial resolution (mm)
| 0.2 | 1.0 | 3.0
Y の値 1 | 0.4 | 0.3 | 6
nGEM
Li-glass
µNID
Boron µNID
MEMS µPIC
µNID
(current)
WCNR-11 4 Sep 2018 J. Parker
LiTA12
(centroiding)

## Slide 30: Summary
Development of detectors to meet the needs of conventional and energy-resolved neutron imaging at RADEN is ongoing
Standard µNID
  Optimized detector components and refined analysis for improved spatial resolution/image quality and improved rate performance
  New analysis GUI and control software for improved ease-of-use
Continuing µNID development
  Basic operation of small-pitch MEMS µPIC confirmed; testing of new glass-substrate MEMS µPIC this winter
  Characterized µNID with boron converter; need to improve efficiency
Optimization of Li-glass detector
  Demonstrated improved spatial resolution with centroiding and sub-pixel image stepping
WCNR-11 4 Sep 2018 J. Parker
Notes: And in summary, we have successfully developed a new, powerful detector system for time-resolved neutron imaging which combines 3-dimensional tracking and energy measurement by time-over-threshold to provide excellent overall operating characteristics.
Two µNID systems are already in service at J-PARC, with a third to be installed at the new imaging beamline.
I have also studied new time-of-flight based techniques and analysis methods, which have applications over a broad range of materials science, and I would like to use such new techniques to open up new science.
I have also developed good working relationships with members of J-PARC and Hokkaido University, and staff at DNP.
I am ready to continue and expand my research at the new imaging beamline, and use the skills I’ve acquired over my varied career to help develop a unique, world class beamline.
Thank you.

## Slide 31: Thank you for your attention!

## Slide 32: Extra slides

## Slide 33: Public beamlines at the MLF
BL02: DNA
BL01: 4 SEASONS
BL11: PLANET
BL15: TAIKAN
BL17: SHARAKU
BL18: SENJU
BL22: RADEN
Mercury spallation target, 25 Hz repetition rate, beam power up to 1 MW
23 beam ports, 6 public beam lines currently in operation
Dedicated imaging beamline at BL22 (RADEN)
WCNR-11 4 Sep 2018 J. Parker

## Slide 34: RADEN/BL22 – Neutron imaging instrument at the MLF
Moderator
Bulk shield
Shutter
Optical Devices
Experimental Space
1st Detector
position
2nd Detector position
13m
5m
Large load sample stage
Optical bench
Medium sample stage
0m
8m
14m
27m
31m
18m
23m
Conventional/pulsed-neutron imaging | Large beam size (up to 30x30 cm2)
High flux (2.6x107 n/s/cm2 @ <0.5eV) | Variable L/D (up to 7500)
Wide bandwidth (~9Å, Δλ/λ < 0.2%) | Large experimental area
Optional sample environments | Automated measurements
Properties of RADEN
WCNR-11 4 Sep 2018 J. Parker

## Slide 35: Performance of standard µNID
Image taken at NOBORU (2011)
Template for fit
Measured TOT distribution
Proton
Triton
137Cs
No source
‘Energy’ cut
Neutrons
γ’s
Total TOT (clocks)
Strong gamma rejection using sum of TOT
Fine spatial resolution from fit to TOT distribution
Performance characteristics |  |  |  |  |  | 
Active area | Spatial resolution | Time resolution | γ-sensitivity | Efficiency @25.3meV | Rate capacity | Effective max. rate
10 x 10 cm2 | 0.1 mm | 0.25 µs | < 10-12 | 26% | 8 Mcps | 1 Mcps
WCNR-11 4 Sep 2018 J. Parker
Notes: The next few slides demonstrate the basic performance of the µNID.
Here are a series of images to show the good spatial resolution we’ve achieved with our detector.
For comparison, the left-most image was constructed by ignoring the TOT information and selecting the mid-point of each track, resulting in a spatial resolution of about 1 mm.
I initially developed a simple algorithm which used the shape of the TOT distribution to determine the proton direction and then apply a correction relative to the mid-point of the track, resulting in the middle image. This method achieves a spatial resolution of about 150 µm but with poor overall image quality.
The final method which I developed uses a template-based fit, with templates generated by a simulation of the detector system. This method achieves a spatial resolution as low as 100 µm with good image uniformity as shown on the right.
The plot in the lower left shows the distribution of total time-over-threshold, summed event-by-event, where we can see a clear separation of the neutron and gamma peaks. By also considering the characteristic track length and TOT distribution for He3 decay events and optimizing the gas gain, we could achieve an effective gamma sensitivity of < 10-12.

## Slide 36: B-µNID and µNID
µNID: 3He converter
Larger event size but better position reconstruction
B-µNID (1.6atm)
Spatial resolution: 0.45 mm
µNID
Spatial resolution: 0.1 mm
WCNR-11 4 Sep 2018 J. Parker

## Slide 37: B-µNID and nGEM
nGEM: similar boron-coating
Ar:CO2 (90:10) at 1 atm
128 x 128 strips, 0.8 mm pitch
Spatial resolution about 1 mm
B-µNID (1.6atm)
Spatial resolution: 0.45 mm
nGEM
Spatial resolution: 1 mm
WCNR-11 4 Sep 2018 J. Parker